# STANDARD OPERATING PROCEDURE

Title: Reverse Response Curve Preparation and Analysis (Experiment 1) for a panel of Phosphopeptide Assays. Version #: 2; Author: Broad Institute Proteomics Platform – Carr Lab. Date: August 1, 2016; BRD-002.

# Purpose

The purpose of this document is to describe the preparation of a reverse response curve for a panel of phosphopeptides.

# Scope

This procedure may be used to prepare a reverse curve of phosphopeptides in a digested background matrix (cells, tissues, plasma, etc) and analyze the mixture after IMAC enrichment by targeted mass spectrometry.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Microcentrifuge

- Quantiva TSQ Mass spectrometer (Thermo)

- Easy NanoLC 1000 liquid chromatograph (Thermo)

# Materials

- HPLC grade water

- formic acid (Fluka, 56302)

- acetonitrile

# Reagents

Standards:

- Heavy stable isotope standard mix (H-SIS): 2 pmol/uL, 30% acetonitrile/0.1% formic acid (New England Peptide)

- Medium stable isotope standard mix (M-SIS): 1.78 pmol/uL, 30% acetonitrile/0.1% formic acid (New England Peptide)

- Michrom mix: 50 fmol/uL digest of 6 equimolar bovine proteins (Michrom Bioresources, #PTD/00001/63) in 3% acetonitrile/5% acetic acid [Note: this product is no longer commercially available]

Matrix:

Equal mixture of two cell line lysates, A375 (melanoma – MAPK) and A549 (lung cancer – KRas), desalted and dried into wells on a 96 well plate in 500 ug aliquots as described in SOP_Broad_SamplePrep&PhosphoEnrich_02, BRD-001, “Cell Lysis, Tryptic Digestion, and Phosphopeptide Enrichment by Automated Immobilized Metal Affinity Chromatography (IMAC)”

# Solutions

- 50% acetonitrile

- 80% acetonitrile

- 3% acetonitrile/5% formic acid

- Reconstitution buffer: 24 uL heavy peptide standard mix, 1176 uL 3% acetonitrile/5% formic acid

# Procedure

- Preparation of Reverse Curve (refer to Table 1)

  - Make standard solution (20 fmol/uL medium stable isotope standard mix (M-SIS) in 80% acetonitrile/0.1% formic acid). Add 14 uL of M-SIS to 1236 uL 80% acetonitrile/0.1% formic acid.

  - Add 112.5 uL standard solution to 8 microcentrifuge tubes (1.5 mL). Cap one tube (tube 1) and set aside to use as blank.

  - Add 37.5 uL H-SIS to tube 8, final concentration 500 fmol/uL.

  - Add 0.4 uL M-SIS to tube 8 to restore medium SIS concentration to 20 fmol/uL after the addition of H-SIS.

  - Transfer 37.5 uL from tube 8 to tube 7 for final heavy SIS concentration of 125 fmol/uL in tube 7. Continue serial dilution, removing 37.5 uL from previous tube and adding to successive tube for a total of 7 tubes.

Table 1. Concentrations of M-SIS and H-SIS in reverse curve

| pt | total M-SIS (fmol) | total H-SIS (fmol) | M-SIS conc (fmol/ug) | H-SIS conc (fmol/ug) |
| --- | --- | --- | --- | --- |
| 1 | 200 | 0 | 0.4 | 0 |
| 2 | 200 | 1.2 | 0.4 | 0.002 |
| 3 | 200 | 5 | 0.4 | 0.01 |
| 4 | 200 | 20 | 0.4 | 0.04 |
| 5 | 200 | 78 | 0.4 | 0.16 |
| 6 | 200 | 312 | 0.4 | 0.63 |
| 7 | 200 | 1250 | 0.4 | 2.5 |
| 8 | 200 | 5000 | 0.4 | 10 |

- Preparation of Sample Plate (refer to Table 2)

  - Add 88 uL 50% acetonitrile to the sample plate containing dried digested desalted cell lysate prepared using SOP_Broad_SamplePrep&PhosphoEnrich_02, BRD-001, “Cell Lysis, Tryptic Digestion, and Phosphopeptide Enrichment by Automated Immobilized Metal Affinity Chromatography (IMAC)”. Refer to the plate map in Table 2.

  - Seal plate well with aluminum foil seal, sonicate and vortex to resuspend peptides. Centrifuge for 30 s at 1500 x g to bring liquid into wells.

  - Remove seal, add 132 uL 100% acetonitrile to make the solution 80% acetonitrile.

  - Seal plate well with a fresh aluminum foil seal, sonicate and vortex to resuspend peptides. Centrifuge for 30 s at 1500 x g to bring liquid into wells. Each well now contains 500 ug cell lysate peptides resuspended in 220 uL 80% acetonitrile.

  - Remove seal, add 10 uL from each curve tube prepared in 1. above and pipet up and down five times to mix.

For example, dispense 10 uL of curve point 1 (0 fmol/uL H-SIS, 20 fmol/uL M-SIS) into wells A1 A2 A3, 10 uL each of curve point 2 (0.02 fmol/uL H-SIS, 20 fmol/uL M-SIS) into wells B1 B2 B3 and continue in a similar manner for each curve point.

  - Follow SOP_Broad_SamplePrep&PhosphoEnrich_02, BRD-001, “Cell Lysis, Tryptic Digestion, and Phosphopeptide Enrichment by Automated Immobilized Metal Affinity Chromatography (IMAC)” to enrich phosphopeptides.

Table 2. Curve platemap and concentrations

|  | plate | map |  |
| --- | --- | --- | --- |
|  | 1 | 2 | 3 |
| A | pt 1-1 | pt 1-2 | pt 1-3 |
| B | pt 2-1 | pt 2-2 | pt 2-3 |
| C | pt 3-1 | pt 3-2 | pt 3-3 |
| D | pt 4-1 | pt 4-2 | pt 4-3 |
| E | pt 5-1 | pt 5-2 | pt 5-3 |
| F | pt 6-1 | pt 6-2 | pt 6-3 |
| G | pt 7-1 | pt 7-2 | pt 7-3 |
| H | pt 8-1 | pt 8-2 | pt 8-3 |

|  | curve | concentration | (fmol/ug) |
| --- | --- | --- | --- |
|  | 1 | 2 | 3 |
| A | 0 | 0 | 0 |
| B | 0.002 | 0.002 | 0.002 |
| C | 0.01 | 0.01 | 0.01 |
| D | 0.04 | 0.04 | 0.04 |
| E | 0.16 | 0.16 | 0.16 |
| F | 0.63 | 0.63 | 0.63 |
| G | 2.5 | 2.5 | 2.5 |
| H | 10 | 10 | 10 |

- Reconstitution of samples for LC-MRM-MS analysis

  - After IMAC enrichment, resuspend dried samples by adding 20 uL of 3% acetonitrile/5% formic acid to each autosampler vial.

  - Sonicate and vortex to resuspend phosphopeptides; centrifuge briefly (20 s at 1000 x g) to bring all liquid to the bottom of autosampler vial.

- NanoLC-MRM-MS analysis

  - set up the TSQ-Quantiva MS with Easy NanoLC1000 and the following:

    - Q1 resolution 0.2

    - Q3 resolution 0.7

    - 1.5 min cycle time

    - CID gas 1.5

    - 0.075 x 100 mm PicoFrit Reprosil C18, 3 um, 200 Å pore size, (Dr. Maisch GmBH) PicoFrit column

    - mobile phase A: 3% acetonitrile/0.1% formic acid, mobile phase B: 90% acetonitrile/0.1% formic acid

    - flowrate: 300 nL/min

    - column temperature: 50 oC

    - gradient: 0 – 35% B in 33 min, 35 – 90% B in 3 min, hold 90% B for 6 min. (42 min total time)

    - injection volume: 2 uL

Prior to analysis, LC-MS/MS of a mixture of H-SIS peptides was performed on a high resolution MS (QExactive) to generate spectral libraries. These were imported into Skyline and used to selected MRM transitions for each peptide. The best 5-10 transition m/z were selected and the H-SIS was analyzed by LC-MRM-MS to determine the optimal collision energy. The best 3-5 transition ions including those required to verify the specificity of the phosphosite were then selected for further sample analysis. See Appendix 2 for scheduled MRM method used to acquire the data.

- Run order

  - Samples are injected by order of increasing concentration, point 1 to point 8, one complete curve followed by the next replicate, with blanks injections and carryover injections performed between each curve. Refer to Table 3 for the complete Run Order.

Table 3. Injection run order of replicate curves

| Curve 1 |  |  | Curve 2 |  |  | Curve 3 |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| inj # | Sample |  | inj # | Sample |  | inj # | Sample |
| 1 | michrom_50fmol_01 |  | 24 | michrom_50fmol_03 |  | 47 | michrom_50fmol_05 |
| 2 | blank_01 |  | 25 | blank_05 |  | 48 | blank_09 |
| 3 | blank_02 |  | 26 | blank_06 |  | 49 | blank_10 |
| 4 | blank_03 |  | 27 | blank_07 |  | 50 | blank_11 |
| 5 | pt1_0fmol_01_01 |  | 28 | pt1_0fmol_02_01 |  | 51 | pt1_0fmol_03_01 |
| 6 | pt1_0fmol_01_02 |  | 29 | pt1_0fmol_02_02 |  | 52 | pt1_0fmol_03_02 |
| 7 | pt1_0fmol_01_03 |  | 30 | pt1_0fmol_02_03 |  | 53 | pt1_0fmol_03_03 |
| 8 | pt1_0fmol_01_04 |  | 31 | pt1_0fmol_02_04 |  | 54 | pt1_0fmol_03_04 |
| 9 | pt1_0fmol_01_05 |  | 32 | pt1_0fmol_02_05 |  | 55 | pt1_0fmol_03_05 |
| 10 | pt2_0002fmol_01_01 |  | 33 | pt2_0002fmol_02_01 |  | 56 | pt2_0002fmol_03_01 |
| 11 | pt3_001fmol_01_01 |  | 34 | pt3_001fmol_02_01 |  | 57 | pt3_001fmol_03_01 |
| 12 | pt4_004fmol_01_01 |  | 35 | pt4_004fmol_02_01 |  | 58 | pt4_004fmol_03_01 |
| 13 | pt5_016fmol_01_01 |  | 36 | pt5_016fmol_02_01 |  | 59 | pt5_016fmol_03_01 |
| 14 | pt6_063fmol_01_01 |  | 37 | pt6_063fmol_02_01 |  | 60 | pt6_063fmol_03_01 |
| 15 | pt7_205fmol_01_01 |  | 38 | pt7_205fmol_02_01 |  | 61 | pt7_205fmol_03_01 |
| 16 | pt8_10fmol_01_01 |  | 39 | pt8_10fmol_02_01 |  | 62 | pt8_10fmol_03_01 |
| 17 | carryover_blank_01 |  | 40 | carryover_blank_03 |  | 63 | carryover_blank_05 |
| 18 | carryover_blank_02 |  | 41 | carryover_blank_04 |  | 64 | carryover_blank_06 |
| 19 | wash_01 |  | 42 | wash_04 |  | 65 | wash_07 |
| 20 | wash_02 |  | 43 | wash_05 |  | 66 | wash_08 |
| 21 | wash_03 |  | 44 | wash_06 |  | 67 | wash_09 |
| 22 | blank_04 |  | 45 | blank_08 |  | 68 | blank_12 |
| 23 | michrom_50fmol_02 |  | 46 | michrom_50fmol_04 |  | 69 | michrom_50fmol_06 |

- Data Analysis

  - Raw files were imported into Skyline.

  - Extracted Ion chromatograms (XIC) of all transition ions were integrated using a Skyline document (Skyline daily version 2.6. https://brendanx-uw1.gs.washington.edu/labkey/project/home/software/Skyline/begin.view).

  - Integrated peaks were manually inspected to confirm proper integration and detection of the transitions for the corresponding light and heavy peptides.

# Referenced Documents

Protocol for automated IMAC enrichment:

Mol Cell Proteomics. 2016 May;15(5):1622-41. doi: 10.1074/mcp.M116.058354. Epub 2016 Feb 24. PMID:26912667

Reduced-representation Phosphosignatures Measured by Quantitative Targeted MS Capture Cellular States and Enable Large-scale Comparison of Drug-induced Phenotypes.

Abelin JG1, Patel J1, Lu X1, Feeney CM1, Fagbami L1, Creech AL1, Hu R1, Lam D1, Davison D1, Pino L1, Qiao JW1, Kuhn E1, Officer A1, Li J2, Abbatiello S1,Subramanian A1, Sidman R2, Snyder E3, Carr SA1, Jaffe JD4

# Appendix 1. Proteins, Peptides and Pathways

| Protein | Pathway | Peptide | Phospho-site |
| --- | --- | --- | --- |
| ARAF | Ras mediated signaling | DSGYYWEVPPSEVQLLK | S299 |
| AURKA | Cell cycle | TTLCGTLDYLPPEMIEGR | T288 |
| AURKB | Cell cycle | KTMCGTLDYLPPEMIEGR | T232 |
| FOXO1 | Signal transduction | TSSNASTISGR | S318 |
| JUN | Immune system | NSDLLTSPDVGLLK | S63 |
| MAPK14 | Signal transduction | HTDDEMTGYVATR | T180 |
| NFKB1 | Signal transduction | DSDSVCDSGVETSFR | S932 |
| PLK1 | Cell cycle | KTLCGTPNYIAPEVLSK | T210 |
| PLK1 | Cell cycle | RSLLELHK | S137 |
| PRKCB | Signal transduction | QPVELTPTDK | T642 |
| RAF1 | Signal transduction | ASDDGKLTDPSK | S43 |
| RAF1 | Signal transduction | SASEPSLHR | S621 |
| RB1 | Cell cycle | TAVIPINGSPR | S249 |
| RICTOR | Signal transduction | AQSLKAPSIATIK | S1302 |
| RPS6 | Disease | LSSLRASTSK | S235 |
| RPS6 | Disease | LSSLRASTSK | S236 |
| RPS6 | Disease | LSSLRASTSK | S240 |
| STAT3 | Signal transduction | FICVTPTTCSNTIDLPMSPR | S727 |
| SUFU | Development pathways | DSLESDSSTAIIPHELIR | S346 |
| TP53BP1 | DNA repair | QYTESQLR | S1778 |

# Appendix 2. Scheduled MRM Method

| Compound | Start Time (min) | End Time (min) | Polarity | Precursor (m/z) | Product (m/z) | Collision Energy (V) |
| --- | --- | --- | --- | --- | --- | --- |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 432.2817 | 17.9 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 632.3978 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 679.3175 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 729.4505 | 20.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 471.9318 | 638.4179 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 471.9318 | 735.4706 | 20.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 440.2959 | 17.9 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 640.412 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 679.3175 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 737.4647 | 20.4 |
| AQS[+80.0]LKAPSIATIK2 | 20.5 | 25 | Positive | 471.9318 | 438.3018 | 17.9 |
| AQS[+80.0]LKAPSIATIK2 | 20.5 | 25 | Positive | 471.9318 | 679.3175 | 15.4 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 371.1197 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 538.2669 | 22.1 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 547.2722 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 660.3563 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 660.2947 | 553.2923 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 660.2947 | 666.3764 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 371.1197 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 542.274 | 22.1 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 555.2864 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 668.3705 | 29.6 |
| AS[+80.0]DDGKLTDPSK2 | 12.5 | 16 | Positive | 660.2947 | 371.1197 | 29.6 |
| AS[+80.0]DDGKLTDPSK2 | 12.5 | 16 | Positive | 660.2947 | 541.277 | 22.1 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 982.3941 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 1010.588 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 1080.371 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 1109.656 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1048.49 | 1016.608 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1048.49 | 1115.677 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 982.3941 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 1018.602 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 1080.371 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 1117.671 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK2 | 31 | 35.5 | Positive | 1048.49 | 982.3941 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK2 | 31 | 35.5 | Positive | 1048.49 | 1080.371 | 27.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 427.1823 | 21.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 764.4413 | 26.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 877.5254 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 990.6095 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 690.3346 | 883.5455 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 690.3346 | 996.6296 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 427.1823 | 21.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 774.4496 | 26.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 887.5337 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 1000.618 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR2 | 26 | 30 | Positive | 690.3346 | 427.1823 | 21.8 |
| DS[+80.0]LESDSSTAIIPHELIR2 | 26 | 30 | Positive | 690.3346 | 770.4615 | 26.8 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 875.3659 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 962.3979 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 1077.425 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 1139.479 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 1237.455 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 885.3741 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 972.4062 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 1087.433 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 1149.487 | 36.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 1247.464 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 881.386 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 968.418 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 1083.445 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 1145.499 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 1243.476 | 29.4 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 446.2722 | 14.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 598.2103 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 609.3355 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 749.3941 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 869.3094 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 456.2804 | 14.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 598.2103 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 619.3438 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 759.4023 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 869.3094 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 452.2923 | 14.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 598.2103 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 615.3556 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 755.4142 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 869.3094 | 19.5 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 672.3927 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 741.3001 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 743.4298 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 952.3958 | 23.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 659.6668 | 678.4128 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 659.6668 | 749.4499 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 680.4069 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 741.3001 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 751.444 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 952.3958 | 23.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK2 | 23.5 | 28 | Positive | 659.6668 | 741.3001 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK2 | 23.5 | 28 | Positive | 659.6668 | 952.3958 | 23.4 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 416.2051 | 30.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 464.7315 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 831.4029 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 928.4557 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 987.3675 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 421.2092 | 30.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 469.7356 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 841.4112 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 938.4639 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 987.3675 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 419.2152 | 30.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 467.7415 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 837.423 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 934.4758 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 987.3675 | 28.1 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 422.2245 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 637.3069 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 762.4468 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 849.4789 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 885.4789 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 430.2387 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 637.3069 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 770.461 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 857.4931 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 885.4789 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 428.2447 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 637.3069 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 768.467 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 855.499 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 422.2245 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 637.3069 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 762.4468 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 929.4452 | 19 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 430.2387 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 637.3069 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 770.461 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 937.4594 | 19 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 428.2447 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 637.3069 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 768.467 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 935.4653 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 502.1909 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 557.3406 | 26.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 842.4132 | 21.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 885.4789 | 31.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 929.4452 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 510.2051 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 557.3406 | 26.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 850.4274 | 21.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 885.4789 | 31.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 937.4594 | 19 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 508.211 | 19 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 557.3406 | 26.5 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 848.4333 | 21.5 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 885.4789 | 31.5 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 935.4653 | 19 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 741.4505 | 28.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 908.4489 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 1009.497 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 1122.581 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 779.397 | 1015.517 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 779.397 | 1128.601 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 749.4647 | 28.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 916.4631 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 1017.511 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 1130.595 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK2 | 28 | 32 | Positive | 779.397 | 747.4706 | 28.7 |
| NSDLLTS[+80.0]PDVGLLK2 | 28 | 32 | Positive | 779.397 | 914.469 | 23.7 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 460.2402 | 25.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 543.2773 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 641.2542 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 754.3383 | 20.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 607.294 | 647.2743 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 607.294 | 760.3584 | 20.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 468.2544 | 25.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 551.2915 | 25.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 649.2684 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 762.3525 | 20.3 |
| QPVELT[+80.0]PTDK2 | 16.5 | 23 | Positive | 607.294 | 466.2603 | 25.3 |
| QPVELT[+80.0]PTDK2 | 16.5 | 23 | Positive | 607.294 | 549.2974 | 17.8 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 359.1935 | 339.2139 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 359.1935 | 452.298 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 359.1935 | 526.2984 | 14.6 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.2002 | 339.2139 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.2002 | 452.298 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.2002 | 532.3185 | 14.6 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.8649 | 339.2139 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.8649 | 452.298 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.8649 | 534.3126 | 14.6 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 532.232 | 357.1405 | 20.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 532.232 | 609.3467 | 25.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 532.232 | 738.3893 | 22.9 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 535.2421 | 357.1405 | 20.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 535.2421 | 615.3668 | 25.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 535.2421 | 744.4094 | 22.9 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 537.2362 | 357.1405 | 20.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 537.2362 | 619.355 | 25.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 537.2362 | 748.3976 | 22.9 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 410.6893 | 20.3 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 722.3944 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 820.3713 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 933.4554 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 605.818 | 826.3914 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 605.818 | 939.4755 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 415.6934 | 20.3 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 732.4027 | 27.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 830.3796 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 943.4636 | 17.8 |
| TAVIPINGS[+80.0]PR2 | 19 | 23 | Positive | 605.818 | 413.6994 | 20.3 |
| TAVIPINGS[+80.0]PR2 | 19 | 23 | Positive | 605.818 | 728.4145 | 17.8 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 533.3042 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 541.1654 | 19.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 620.3362 | 24.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 691.3733 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 583.7609 | 626.3563 | 24.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 583.7609 | 697.3935 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 541.1654 | 19.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 543.3125 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 630.3445 | 24.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 701.3816 | 27 |
| TSS[+80.0]NASTISGR2 | 11 | 15 | Positive | 583.7609 | 539.3243 | 27 |
| TSS[+80.0]NASTISGR2 | 11 | 15 | Positive | 583.7609 | 541.1654 | 19.5 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 416.2051 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 464.7315 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 616.2759 | 22.6 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 831.4029 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 928.4557 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 421.2092 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 469.7356 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 616.2759 | 22.6 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 841.4112 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 938.4639 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 419.2152 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 467.7415 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 616.2759 | 22.6 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 837.423 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 934.4758 | 20.1 |