# STANDARD OPERATING PROCEDURE

Title: Repeatability/Validation (Experiment 2) for a panel of phosphopeptide assays. Version #: 2; Author: Broad Institute Proteomics Platform – Carr Lab. Date: August 1, 2016; BRD-003.

# Purpose

The purpose of this document is to describe the generation of a repeatability/validation experiment for a panel of phosphopeptides.

# Scope

This procedure may be used to evaluate the repeatability/validation of IMAC enrichment and analysis by targeted mass spectrometry of phosphopeptides in a digested background matrix (cells, tissues, plasma, etc).

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Microcentrifuge

- Quantiva TSQ Mass spectrometer (Thermo)

- Easy NanoLC 1000 liquid chromatograph (Thermo)

# Materials

- HPLC water

- formic acid (Fluka, 56302)

- acetonitrile

# Reagents

Standards:

- Heavy stable isotope standard mix (H-SIS): 2 pmol/uL, 30% acetonitrile/0.1% formic acid (New England Peptide)

- Medium stable isotope standard mix (M-SIS): 1.78 pmol/uL, 30% acetonitrile/0.1% formic acid (New England Peptide)

- Michrom mix: 50 fmol/uL digest of 6 equimolar bovine proteins (Michrom Bioresources, #PTD/00001/63) in 3% acetonitrile/5% acetic acid [Note: this product is no longer commercially available]

Matrix:

Equal mixture of two cell line lysates, A375 (melanoma – MAPK) and A549 (lung cancer – KRas), desalted and dried into wells on a 96 well plate in 500 ug aliquots as described in SOP_Broad_SamplePrep&PhosphoEnrich_02, BRD-001, “Cell Lysis, Tryptic Digestion, and Phosphopeptide Enrichment by Automated Immobilized Metal Affinity Chromatography (IMAC)”

# Solutions

- 50% acetonitrile

- 80% acetonitrile

- 3% acetonitrile/5% formic acid

- Reconstitution buffer: 24 uL heavy peptide standard mix, 1176 uL 3% acetonitrile/5% formic acid

# Procedure

- Determination of spike levels

  - Following the guidelines for Experiment 2 as described in the CPTAC Assay Development Guidelines, three levels of concentration, “Lo” , “Med” and “Hi” were selected for repeatability experiments based on the reverse curve performance\.

    - “Lo” = 0.16 fmol/ug

    - “Med” = 0.48 fmol/ug

    - “Hi” = 1.92 fmol/ug

  - Five replicates of these samples were prepared by adding H-SIS phosphopeptides into digested, desalted lysate at each of the three concentration levels. These samples were then frozen. On each of 5 non-overlapping consecutive days, a set of samples were thawed and spiked with standard amount of M-SIS phosphopeptides (five replicates of each M-SIS were also prepared and frozen), enriched using automated IMAC and analyzed by targeted mass spectrometry.

- Preparation of Repeatability Samples - Completed before any enrichments

  - Cell lysates were prepared as described in SOP_Broad_SamplePrep&PhosphoEnrich_02, BRD-001, “Cell Lysis, Tryptic Digestion, and Phosphopeptide Enrichment by Automated Immobilized Metal Affinity Chromatography (IMAC)”.

  - Resuspend cell lysates in 50% acetonitrile to a final concentration of 2.5 mg/ml.

  - Label 15 microcentrifuge tubes (1.5 mL):  label five “Lo”, five “Med” and five “Hi”.

  - Aliquot 660 uL of 2.5 mg/mL digested, desalted lysate to each tube.

  - Prepare a mixture of heavy phosphospeptides from H-SIS stock mixture (2 pmol/uL) in a single tube for each of the 3 levels of concentration as described below:

    - Lo – (0.16 fmol/ug): 2.4 uL H-SIS and 97.6 uL 80% ACN

    - Med – (0.48 fmol/ug): 7.2 uL H-SIS and 92.8 uL 80% ACN

    - Hi – (1.92 fmol/ug): 28.8 uL H-SIS and 71.2 uL 80% ACN

  - Prepare five tubes for each concentration of Lo, Med, and Hi from the working stock mixtures in a single tube for each of the 3 levels of concentration as described below:

    - Add 5 uL of the Lo (0.16 fmol/ug) to each tube labeled “Lo”.

    - Add 5 uL of the Med (0.48 fmol/ug) to each tube labeled “Med”.

    - Add 5 uL of the Hi (1.92 fmol/ug) to each tube labeled “Hi”.

  - Freeze all 15 sample tubes at -80oC.

  - Aliquot 7 uL M-SIS phosphopeptide stock (1.78 pmol/uL) into 5 tubes labeled “Med SIS” and freeze these aliquots at -80oC.

- Preparation of Sample Plate – Performed each day of enrichment for 5 days

  - Remove a single aliquot of each of the following previously prepped samples:

    - “Med SIS” stock (7 uL @ 1.78 pmol/uL)

    - “Hi” sample

    - “Med” sample

    - “Lo” sample

  - Add 93 uL 80% ACN to “Med SIS” stock tube to bring final concentration to 120 fmol/uL.

  - Add 5 uL of 120 fmol/uL Med SIS to each of the 3 sample tubes (“Hi”, “Med”, “Lo”).

  - Divide each concentration point into process triplicate on the plate, as depicted below:

  - Follow IMAC enrichment protocol on AssayMAP Bravo as described in SOP_Broad_SamplePrep&PhosphoEnrich_02, BRD-001, “Cell Lysis, Tryptic Digestion, and Phosphopeptide Enrichment by Automated Immobilized Metal Affinity Chromatography (IMAC)” to enrich phosphopeptides.

- Reconstitution of samples for analysis by mass spectrometry

  - After IMAC enrichment samples will be dried in autosampler vials.

  - Resuspend dried samples in 20 uL of Reconstitution buffer (24 uL heavy peptide standard mix, 1176 uL 3% acetonitrile/5% formic acid, prepared fresh).

  - Sonicate and vortex to resuspend phosphopeptides; centrifuge briefly (20 s at 1000 x g) to bring all liquid to the bottom of autosampler vial.

- NanoLC-MRM-MS analysis

  - set up the TSQ-Quantiva MS with Easy NanoLC1000 and the following:

    - Q1 resolution 0.2

    - Q3 resolution 0.7

    - 1.5 min cycle time

    - CID gas 1.5

    - 0.075 x 100 mm PicoFrit Reprosil C18, 3 um, 200 Å pore size, (Dr. Maisch GmBH) PicoFrit column

    - mobile phase A: 3% acetonitrile/0.1% formic acid, mobile phase B: 90% acetonitrile/0.1% formic acid

    - flowrate: 300 nL/min

    - column temperature: 50 oC

    - gradient: 0 – 35% B in 33 min, 35 – 90%B in 3 min, hold 90%B for 6 min. (42 min total time)

    - injection volume: 2 uL

  - Prior to analysis, LC-MS/MS of a mixture of H-SIS peptides was performed on a high resolution MS (QExactive) to generate spectral libraries. These were imported into Skyline and used to selected MRM transitions for each peptide. The best 5-10 transition m/z’s were selected and the H-SIS was analyzed by LC-MRM-MS to determine the optimal collision energy. The best 3-5 transition ions including those required to verify the specificity of the phosphosite were then selected for further sample analysis. See Appendix 2 for scheduled MRM method used to acquire the data.

- Run order

  - Samples are analyzed in random run order as described below:

- Data Analysis

  - Raw files were imported into Skyline.

  - Extracted Ion chromatograms (XIC) of all transition ions were integrated using a Skyline document (Skyline daily version 2.6. https://brendanx-uw1.gs.washington.edu/labkey/project/home/software/Skyline/begin.view).

  - Integrated peaks were manually inspected to confirm proper integration and detection of the transitions for the corresponding light and heavy peptides.

# Referenced Documents

Protocol for automated IMAC enrichment:

Mol Cell Proteomics. 2016 May;15(5):1622-41. doi: 10.1074/mcp.M116.058354. Epub 2016 Feb 24. PMID:26912667

Reduced-representation Phosphosignatures Measured by Quantitative Targeted MS Capture Cellular States and Enable Large-scale Comparison of Drug-induced Phenotypes.

Abelin JG1, Patel J1, Lu X1, Feeney CM1, Fagbami L1, Creech AL1, Hu R1, Lam D1, Davison D1, Pino L1, Qiao JW1, Kuhn E1, Officer A1, Li J2, Abbatiello S1,Subramanian A1, Sidman R2, Snyder E3, Carr SA1, Jaffe JD4

# Appendix 1. Proteins, Peptides and Pathways

| Protein | Pathway | Peptide | Phospho-site |
| --- | --- | --- | --- |
| ARAF | Ras mediated signaling | DSGYYWEVPPSEVQLLK | S299 |
| AURKA | Cell cycle | TTLCGTLDYLPPEMIEGR | T288 |
| AURKB | Cell cycle | KTMCGTLDYLPPEMIEGR | T232 |
| FOXO1 | Signal transduction | TSSNASTISGR | S318 |
| JUN | Immune system | NSDLLTSPDVGLLK | S63 |
| MAPK14 | Signal transduction | HTDDEMTGYVATR | T180 |
| NFKB1 | Signal transduction | DSDSVCDSGVETSFR | S932 |
| PLK1 | Cell cycle | KTLCGTPNYIAPEVLSK | T210 |
| PLK1 | Cell cycle | RSLLELHK | S137 |
| PRKCB | Signal transduction | QPVELTPTDK | T642 |
| RAF1 | Signal transduction | ASDDGKLTDPSK | S43 |
| RAF1 | Signal transduction | SASEPSLHR | S621 |
| RB1 | Cell cycle | TAVIPINGSPR | S249 |
| RICTOR | Signal transduction | AQSLKAPSIATIK | S1302 |
| RPS6 | Disease | LSSLRASTSK | S235 |
| RPS6 | Disease | LSSLRASTSK | S236 |
| RPS6 | Disease | LSSLRASTSK | S240 |
| STAT3 | Signal transduction | FICVTPTTCSNTIDLPMSPR | S727 |
| SUFU | Development pathways | DSLESDSSTAIIPHELIR | S346 |
| TP53BP1 | DNA repair | QYTESQLR | S1778 |

# Appendix 2. Scheduled MRM Method

| Compound | Start Time (min) | End Time (min) | Polarity | Precursor (m/z) | Product (m/z) | Collision Energy (V) |
| --- | --- | --- | --- | --- | --- | --- |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 432.2817 | 17.9 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 632.3978 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 679.3175 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 469.9251 | 729.4505 | 20.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 471.9318 | 638.4179 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 471.9318 | 735.4706 | 20.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 440.2959 | 17.9 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 640.412 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 679.3175 | 15.4 |
| AQS[+80.0]LKAPSIATIK | 20.5 | 25 | Positive | 472.5965 | 737.4647 | 20.4 |
| AQS[+80.0]LKAPSIATIK2 | 20.5 | 25 | Positive | 471.9318 | 438.3018 | 17.9 |
| AQS[+80.0]LKAPSIATIK2 | 20.5 | 25 | Positive | 471.9318 | 679.3175 | 15.4 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 371.1197 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 538.2669 | 22.1 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 547.2722 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 657.2847 | 660.3563 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 660.2947 | 553.2923 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 660.2947 | 666.3764 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 371.1197 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 542.274 | 22.1 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 555.2864 | 29.6 |
| AS[+80.0]DDGKLTDPSK | 12.5 | 16 | Positive | 661.2918 | 668.3705 | 29.6 |
| AS[+80.0]DDGKLTDPSK2 | 12.5 | 16 | Positive | 660.2947 | 371.1197 | 29.6 |
| AS[+80.0]DDGKLTDPSK2 | 12.5 | 16 | Positive | 660.2947 | 541.277 | 22.1 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 982.3941 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 1010.588 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 1080.371 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1045.48 | 1109.656 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1048.49 | 1016.608 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1048.49 | 1115.677 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 982.3941 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 1018.602 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 1080.371 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK | 31 | 35.5 | Positive | 1049.487 | 1117.671 | 32.8 |
| DS[+80.0]GYYWEVPPSEVQLLK2 | 31 | 35.5 | Positive | 1048.49 | 982.3941 | 27.8 |
| DS[+80.0]GYYWEVPPSEVQLLK2 | 31 | 35.5 | Positive | 1048.49 | 1080.371 | 27.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 427.1823 | 21.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 764.4413 | 26.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 877.5254 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 688.3279 | 990.6095 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 690.3346 | 883.5455 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 690.3346 | 996.6296 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 427.1823 | 21.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 774.4496 | 26.8 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 887.5337 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR | 26 | 30 | Positive | 691.664 | 1000.618 | 29.3 |
| DS[+80.0]LESDSSTAIIPHELIR2 | 26 | 30 | Positive | 690.3346 | 427.1823 | 21.8 |
| DS[+80.0]LESDSSTAIIPHELIR2 | 26 | 30 | Positive | 690.3346 | 770.4615 | 26.8 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 875.3659 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 962.3979 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 1077.425 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 1139.479 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 870.8246 | 1237.455 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 885.3741 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 972.4062 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 1087.433 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 1149.487 | 36.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR | 19 | 23 | Positive | 875.8287 | 1247.464 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 881.386 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 968.418 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 1083.445 | 26.9 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 1145.499 | 29.4 |
| DSDSVC[+57.0]DSGVETS[+80.0]FR2 | 19 | 23 | Positive | 873.8346 | 1243.476 | 29.4 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 446.2722 | 14.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 598.2103 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 609.3355 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 749.3941 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 525.8764 | 869.3094 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 456.2804 | 14.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 598.2103 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 619.3438 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 759.4023 | 19.5 |
| HTDDEMT[+80.0]GYVATR | 17 | 21 | Positive | 529.2124 | 869.3094 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 452.2923 | 14.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 598.2103 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 615.3556 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 755.4142 | 19.5 |
| HTDDEMT[+80.0]GYVATR2 | 17 | 21 | Positive | 527.8831 | 869.3094 | 19.5 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 672.3927 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 741.3001 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 743.4298 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 657.6601 | 952.3958 | 23.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 659.6668 | 678.4128 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 659.6668 | 749.4499 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 680.4069 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 741.3001 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 751.444 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK | 23.5 | 28 | Positive | 660.3315 | 952.3958 | 23.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK2 | 23.5 | 28 | Positive | 659.6668 | 741.3001 | 18.4 |
| KT[+80.0]LC[+57.0]GTPNYIAPEVLSK2 | 23.5 | 28 | Positive | 659.6668 | 952.3958 | 23.4 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 416.2051 | 30.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 464.7315 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 831.4029 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 928.4557 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 730.9926 | 987.3675 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 421.2092 | 30.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 469.7356 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 841.4112 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 938.4639 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR | 28 | 32 | Positive | 734.3287 | 987.3675 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 419.2152 | 30.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 467.7415 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 837.423 | 28.1 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 934.4758 | 20.6 |
| KT[+80.0]MC[+57.0]GTLDYLPPEMIEGR2 | 28 | 32 | Positive | 732.9993 | 987.3675 | 28.1 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 422.2245 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 637.3069 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 762.4468 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 849.4789 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 565.2843 | 885.4789 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 430.2387 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 637.3069 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 770.461 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 857.4931 | 31.5 |
| LS[+80.0]SLRASTSK | 12 | 17 | Positive | 569.2914 | 885.4789 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 428.2447 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 637.3069 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 768.467 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 855.499 | 31.5 |
| LS[+80.0]SLRASTSK2 | 12 | 17 | Positive | 568.2943 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 422.2245 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 637.3069 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 762.4468 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 565.2843 | 929.4452 | 19 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 430.2387 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 637.3069 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 770.461 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK | 12 | 17 | Positive | 569.2914 | 937.4594 | 19 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 428.2447 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 637.3069 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 768.467 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 885.4789 | 31.5 |
| LSS[+80.0]LRASTSK2 | 12 | 17 | Positive | 568.2943 | 935.4653 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 502.1909 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 557.3406 | 26.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 842.4132 | 21.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 885.4789 | 31.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 565.2843 | 929.4452 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 510.2051 | 19 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 557.3406 | 26.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 850.4274 | 21.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 885.4789 | 31.5 |
| LSSLRAS[+80.0]TSK | 12 | 17 | Positive | 569.2914 | 937.4594 | 19 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 508.211 | 19 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 557.3406 | 26.5 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 848.4333 | 21.5 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 885.4789 | 31.5 |
| LSSLRAS[+80.0]TSK2 | 12 | 17 | Positive | 568.2943 | 935.4653 | 19 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 741.4505 | 28.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 908.4489 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 1009.497 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 776.3869 | 1122.581 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 779.397 | 1015.517 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 779.397 | 1128.601 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 749.4647 | 28.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 916.4631 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 1017.511 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK | 28 | 32 | Positive | 780.394 | 1130.595 | 23.7 |
| NSDLLTS[+80.0]PDVGLLK2 | 28 | 32 | Positive | 779.397 | 747.4706 | 28.7 |
| NSDLLTS[+80.0]PDVGLLK2 | 28 | 32 | Positive | 779.397 | 914.469 | 23.7 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 460.2402 | 25.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 543.2773 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 641.2542 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 604.2839 | 754.3383 | 20.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 607.294 | 647.2743 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 607.294 | 760.3584 | 20.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 468.2544 | 25.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 551.2915 | 25.3 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 649.2684 | 17.8 |
| QPVELT[+80.0]PTDK | 16.5 | 23 | Positive | 608.291 | 762.3525 | 20.3 |
| QPVELT[+80.0]PTDK2 | 16.5 | 23 | Positive | 607.294 | 466.2603 | 25.3 |
| QPVELT[+80.0]PTDK2 | 16.5 | 23 | Positive | 607.294 | 549.2974 | 17.8 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 359.1935 | 339.2139 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 359.1935 | 452.298 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 359.1935 | 526.2984 | 14.6 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.2002 | 339.2139 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.2002 | 452.298 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.2002 | 532.3185 | 14.6 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.8649 | 339.2139 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.8649 | 452.298 | 17.1 |
| RS[+80.0]LLELHK | 17 | 21 | Positive | 361.8649 | 534.3126 | 14.6 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 532.232 | 357.1405 | 20.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 532.232 | 609.3467 | 25.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 532.232 | 738.3893 | 22.9 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 535.2421 | 357.1405 | 20.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 535.2421 | 615.3668 | 25.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 535.2421 | 744.4094 | 22.9 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 537.2362 | 357.1405 | 20.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 537.2362 | 619.355 | 25.4 |
| SAS[+80.0]EPSLHR | 12 | 16.5 | Positive | 537.2362 | 748.3976 | 22.9 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 410.6893 | 20.3 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 722.3944 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 820.3713 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 602.8079 | 933.4554 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 605.818 | 826.3914 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 605.818 | 939.4755 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 415.6934 | 20.3 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 732.4027 | 27.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 830.3796 | 17.8 |
| TAVIPINGS[+80.0]PR | 19 | 23 | Positive | 607.8121 | 943.4636 | 17.8 |
| TAVIPINGS[+80.0]PR2 | 19 | 23 | Positive | 605.818 | 413.6994 | 20.3 |
| TAVIPINGS[+80.0]PR2 | 19 | 23 | Positive | 605.818 | 728.4145 | 17.8 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 533.3042 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 541.1654 | 19.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 620.3362 | 24.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 580.7508 | 691.3733 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 583.7609 | 626.3563 | 24.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 583.7609 | 697.3935 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 541.1654 | 19.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 543.3125 | 27 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 630.3445 | 24.5 |
| TSS[+80.0]NASTISGR | 11 | 15 | Positive | 585.7549 | 701.3816 | 27 |
| TSS[+80.0]NASTISGR2 | 11 | 15 | Positive | 583.7609 | 539.3243 | 27 |
| TSS[+80.0]NASTISGR2 | 11 | 15 | Positive | 583.7609 | 541.1654 | 19.5 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 416.2051 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 464.7315 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 616.2759 | 22.6 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 831.4029 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 715.9914 | 928.4557 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 421.2092 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 469.7356 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 616.2759 | 22.6 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 841.4112 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR | 31 | 35.5 | Positive | 719.3275 | 938.4639 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 419.2152 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 467.7415 | 20.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 616.2759 | 22.6 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 837.423 | 30.1 |
| TT[+80.0]LC[+57.0]GTLDYLPPEMIEGR2 | 31 | 35.5 | Positive | 717.9981 | 934.4758 | 20.1 |