# STANDARD OPERATING PROCEDURE

Title: Liquid Chromatography using Dionex UltiMate 3000 RSLCnano LC system for PRM assays. Version #: 1.1; Author: Hui Zhang Laboratory – Johns Hopkins University. Date: 02/18/2016.

# Purpose

The purpose of this document is to describe the liquid chromatography (LC) method for quantitative mass spectrometry-based analyses.

# Scope

This procedure describes the setup of the LC and the method parameters. It is specific to the operation of the Dionex UltiMate 3000 RSLCnano LC system.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- HPLC: Dionex UltiMate 3000 RSLCnano LC system (Thermo Fisher Scientific; cat. # NCS-3500RS)

# Materials

- Injection loop: 20 µL Dionex nanoViper sample loop (Thermo Fisher Scientific; cat. # 6826.2420)

- Column compartment: UltiMate 3000 Binary Rapid Separation Nano Flow Pump with Ternary Loading Pump and Column Compartment (Thermo Fisher Scientific; cat. # NCS-3500RS)

- Autosampler: UltiMate 3000 Thermostatted Rapid Separation Pulled Loop Wellplate Sampler (Thermo Fisher Scientific; cat. # WPS-3000)

- Solvent Degasser: UltiMate 3000 Integrated Solvent and Degasser Rack, 4 Channels (Thermo Fisher Scientific; cat. # SRD-3400)

- Trap Column: 300 µm I.D. x 5 mm packed with Acclaim PepMap 100 5 µm, 100 Å C18 (Thermo Fisher Scientific; cat. # 160454)

- Analytical Column: 75 µm I.D. x 50 cm EASY-Spray column packed with Acclaim PepMap RSLC C18, 2 µm (Thermo Fisher Scientific; cat. # ES803)

- Autosampler vials: 9 mm assembled amber autosampler vial kit (Thermo Fisher Scientific; cat. # C5000-196W)

- Autosampler vial inserts: Polyspring, glass conical insert (Thermo Fisher Scientific; cat. # C4010-630)

# Reagents

- Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)

- Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)

- Formic Acid: LC-MS Ultra (Sigma-Aldrich; cat. # 14265)

# Solutions

- Loading pump, mobile phase A: 2% ACN/0.1% formic acid in water

- NanoFlow pump, mobile phase A: 2% ACN/0.1% formic acid in water

- NanoFlow pump, mobile phase B: 0.1% formic acid in 90% ACN

# Procedure

- Autosampler method

Temperature=                    5°C

Draw Speed =                     200 [nl/s]

Draw Delay =                      5000 [ms]

Disp Speed =                       2000 [nl/s]

Dispense Delay =               2000 [ms]

Waste Speed =                    4000 [nl/s]

Wash Speed =                     4000 [nl/s]

Loop Wash Factor =         2.000

Sample Height =                2.000 [mm]

Puncture Depth =              8.000 [mm]

Wash Volume =                 150.000 [µl]

Rinse Between Reinjections = No

Low Dispersion Mode =            Off

Inject Mode =                            µL pick-up

Inject Volume = 	             5.000 µL

- Loading pump profile

  - Flow rate: 5.000 µL/min

  - Flow gradient: Isocratic; 100% A

  - Injection valve switch: 6.000 min

- Gradient method:

  - Flow rate: 250 nL/min

  - Temperature: 42 °C

  - Timetable

| Retention (min) | Flow (µL/min) | %B |
| --- | --- | --- |
| 0.00 | 0.250 | 4.00 |
| 7.00 | 0.250 | 4.00 |
| 27.00 | 0.250 | 40.00 |
| 30.00 | 0.250 | 90.00 |
| 37.00 | 0.250 | 90.00 |
| 38.00 | 0.250 | 4.00 |
| 45.00 | 0.250 | 4.00 |

# Referenced Documents

1. Thomas SN, Harlan R, Chen J, Aiyetan P, Liu Y, Sokoll LJ, Aebersold R, Chan DW, Zhang H. Multiplexed Targeted Mass Spectrometry-Based Assays for the Quantification of N-Linked Glycosite-Containing Peptides in Serum. Anal Chem. 2015 Nov 3;87(21):10830-8. doi: 10.1021/acs.analchem.5b02063. Epub 2015 Oct 21. PubMed PMID: 26451657; PubMed Central PMCID: PMC4708883.