# STANDARD OPERATING PROCEDURE

Title: Parallel Reaction Monitoring (PRM) using a Q-Exactive mass spectrometer. Version #: 1.1; Author: Hui Zhang Laboratory – Johns Hopkins University. Date: 02/18/2016.

# Purpose

The purpose of this document is to describe the mass spectrometry method for the quantitative analysis of peptides using parallel reaction monitoring (PRM).

# Scope

This procedure describes the setup of the mass spectrometer and the PRM method parameters for the Q-Exactive. LC parameters are contained in a separate document.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Source: EASY-Spray (Thermo Fisher Scientific)

- Emitter tip: Precision positioned glass emitter integrated into EASY-Spray column (Thermo Fisher Scientific)

- LC-to-source connection: nanoViper fitting; integrated zero dead volume union within EASY-Spray column (Thermo Fisher Scientific)

- Mass spectrometer: Q-Exactive (Thermo Fisher Scientific)

# Materials

- Injection loop: 20 µL Dionex nanoViper sample loop (Thermo Fisher Scientific; cat. # 6826.2420)

- Column compartment: UltiMate 3000 Binary Rapid Separation Nano Flow Pump with Ternary Loading Pump and Column Compartment (Thermo Fisher Scientific; cat. # NCS-3500RS)

- Autosampler: UltiMate 3000 Thermostatted Rapid Separation Pulled Loop Wellplate Sampler (Thermo Fisher Scientific; cat. # WPS-3000)

- Solvent Degasser: UltiMate 3000 Integrated Solvent and Degasser Rack, 4 Channels (Thermo Fisher Scientific; cat. # SRD-3400)

- Trap Column: 300 µm I.D. x 5 mm packed with Acclaim PepMap 100 5 µm, 100 Å C18 (Thermo Fisher Scientific; cat. # 160454)

- Analytical Column: 75 µm I.D. x 50 cm EASY-Spray column packed with Acclaim PepMap RSLC C18, 2 µm (Thermo Fisher Scientific; cat. # ES803)

# Reagents

- Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)

- Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)

- Formic Acid: LC-MS Ultra (Sigma-Aldrich; cat. # 14265)

# Solutions

- Loading pump, mobile phase A: 2% ACN/0.1% formic acid in water

- NanoFlow pump, mobile phase A: 2% ACN/0.1% formic acid in water

- NanoFlow pump, mobile phase B: 0.1% formic acid in 90% ACN

# Procedure

- Setup MS method and tune file parameters

  - NSI Source/Gas parameters

    - Spray Voltage: 1800 V

    - Capillary Temperature: 250 °C

    - Sheath Gas: 0

    - Auxiliary Gas: 0

  - Scheduled PRM parameters

    - PRM detection window (sec): 300

    - Maximum IT: 100 ms

    - Inclusion list: On

  - MS parameters

    - Microscans: 1

    - Resolution: 70,000

    - AGC target: 5e5

    - Isolation window: 2.0 m/z

    - NCE: 29

    - Lock masses: 445.12003 m/z; 371.10123 m/z

- Test system suitability with 900 ng bovine fetuin tryptic digest once column is conditioned

- Identify scheduling time for target peptides/transitions

  - LC-PRM method preparation

    - Load the Skyline file containing peptides and transitions that will be monitored during the QC analysis

    - Note: For the Q-Exactive, Thermo does not use a true collision offset in their instrument method editor; thus, there is no Q-Exactive-specific collision energy equation in Skyline

    - Export the unscheduled transition list as a single method

    - Import the unscheduled transition list as an Inclusion list in a Targeted-MS2  acquisition method in Xcalibur with the parameters set as indicated in Step 1, above

  - Timing the peptide detection

    - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

    - Inject the QC sample 3x

    - Import the data files into the Skyline file

    - Check the automatic integration of all peaks

    - Export the scheduled transition list using 300 sec scheduled windows

  - QC sample analysis

    - Import the scheduled transition list into an Inclusion list in a Targeted-MS2  acquisition method in Xcalibur with the parameters set as indicated in Step 1, above

    - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

    - Inject the QC sample 3x

# Referenced Documents

1. Thomas SN, Harlan R, Chen J, Aiyetan P, Liu Y, Sokoll LJ, Aebersold R, Chan DW, Zhang H. Multiplexed Targeted Mass Spectrometry-Based Assays for the Quantification of N-Linked Glycosite-Containing Peptides in Serum. Anal Chem. 2015 Nov 3;87(21):10830-8. doi: 10.1021/acs.analchem.5b02063. Epub 2015 Oct 21. PubMed PMID: 26451657; PubMed Central PMCID: PMC4708883.