STANDARD OPERATING PROCEDURE

Title: Multiple Reaction Monitoring (MRM) using a TSQ Vantage triple quadrupole mass spectrometer. Version #: 1.1; Author: Hui Zhang Laboratory – Johns Hopkins University. Date: 06/10/2016.

# Purpose

The purpose of this document is to describe the mass spectrometry method for the quantitative analysis of peptides using multiple reaction monitoring (MRM).

# Scope

This procedure describes the setup of the mass spectrometer and the MRM method parameters for the TSQ Vantage triple quadrupole mass spectrometer. LC parameters are contained in a separate document.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Mass spectrometer: TSQ Vantage (Thermo Scientific)

- Ion source: HESI (Thermo Scientific)

# Procedure

- Setup MS method and tune file parameters

  - ESI Source/Gas parameters

    - Spray voltage: 4000 V

    - Capillary temperature: 210 °C

    - Sheath gas pressure: 15

    - Collision gas pressure: 1.5 mTorr

  - MS parameters

    - Q1 & Q3 peak width (FWHM): 0.70

    - Cycle time: 0.480 sec

    - Declustering potential: 2 V

    - Chrom filter peak width: 30.0 sec

    - Scheduled window: 180 sec

    - Signal trigger: 1.0 e5

- Test system suitability with 100 fmol 6 protein digest standard, equimolar (Thermo Scientific; cat. # 88342) – QC sample

- LC-MRM method preparation and execution

  - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

  - Inject the QC sample three times to assess LC-MRM-MS performance

  - Create Skyline file containing .fasta files of assay targets

  - Export unscheduled transition lists (≤ 60 transitions per list)

  - Run heavy target peptides using unscheduled methods and obtain retention times

  - Import .raw files into Skyline and refine transitions

  - Add light peptides into Skyline document

  - Export a scheduled MRM-MS transition list with a 3 min retention time window

  - Run the scheduled method twice

  - Import the raw files into Skyline and check the peak integration

  - Run QC sample at least once every 16 hours