STANDARD OPERATING PROCEDURE

Title: Liquid Chromatography using Accela system for MRM assays. Version #: 1.1; Author: Hui Zhang Laboratory – Johns Hopkins University. Date: 06/10/2016.

# Purpose

The purpose of this document is to describe the liquid chromatography (LC) method for quantitative mass spectrometry-based analyses.

# Scope

This procedure describes the setup of the LC and the method parameters. It is specific to the operation of the Thermo Scientific Accela system.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- HPLC: Accela 1250 Quaternary Low Pump (Thermo Scientific)

- Autosampler: Accela Open Autosampler (Thermo Scientific)

# Materials

- Injection loop: 10 µL

- Column heater: Hot pocket (Thermo Scientific; cat. # 92016-150)

- Analytical Column: 1.0 mm I.D. x 15 cm Viva C18 3 µm 300 Å (Restek; cat. # 9514361)

- Autosampler vials: Polypropylene 12 x 32 mm screw neck vial with caps and PTFE/silicone septum, 300 µL (Waters; cat. # 186002640)

- Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)

- Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)

- Formic Acid: LC-MS Ultra (Sigma-Aldrich; cat. # 14265)

# Solutions

- Gradient pump, mobile phase A: Water/0.2 % formic acid

- Gradient pump, mobile phase B: 90% ACN/0.2% formic acid

# Procedure

- Autosampler method

Temperature:		4°C

Syringe:		100 µL

Airgap Volume (µl): 		3

Front Volume (µl): 		2

Rear Volume (µl): 	2

Filling Speed (µl/s):	 5

Pullup Delay (ms): 	3000

Inject to: 		LC Vlv1

Injection Speed (µl/s):	 5

Pre & Post Inject Delay (ms): 	500, 500

Needle Gap Valve Clean (mm): 3

Valve Clean Time Solvent 2 (s): 2

Post Clean Time Solvent 2 (s): 2

Valve Clean Time Solvent 1 (s): 5

Post Clean Time Solvent 1 (s): 5

Stator Wash: 		0

Delay Stator Wash (s): 	120

Stator Wash Time Solvent 2 (s): 5

Stator Wash Time Solvent 1 (s): 5

- Gradient method:

  - Flow rate: 50 µL/min

  - Column heater temperature: 50 °C

  - Timetable:

| Retention (min) | Flow (µL/min) | %B |
| --- | --- | --- |
| 0.00 | 50 | 1.00 |
| 5.00 | 50 | 1.00 |
| 30.00 | 50 | 35.00 |
| 35.00 | 50 | 55.00 |
| 38.00 | 50 | 95.00 |
| 45.00 | 50 | 95.00 |
| 46.00 | 50 | 1.00 |
| 60.00 | 50 | 1.00 |