STANDARD OPERATING PROCEDURE

Title: Multiple Reaction Monitoring (MRM) using an 8040 triple quadrupole mass spectrometer. Version #: 1.1; Author: Hui Zhang Laboratory – Johns Hopkins University. Date: 06/10/2016.

# Purpose

The purpose of this document is to describe the mass spectrometry method for the quantitative analysis of peptides using multiple reaction monitoring (MRM).

# Scope

This procedure describes the setup of the mass spectrometer and the MRM method parameters for the 8040 triple quadrupole mass spectrometer. LC parameters are contained in a separate document.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Mass spectrometer: LCMS 8040 (Shimadzu Scientific)

# Procedure

- Setup MS method and tune file parameters

  - ESI Source/Gas parameters

    - Spray voltage: 4500 V

    - Desolvation line temperature: 210 °C

    - Nebulizing gas flow rate: 3 L/min

    - Drying gas flow rate: 15 L/min

    - Heat block temperature: 400 °C

  - MS parameters

    - Q1 & Q3 resolution: Unit

    - Detection window: 180 sec

    - Maximum loop time: 1.59 sec

    - Maximum dwell time: 50 msec

- Test system suitability with 500 fmol E. coli Beta galactosidase tryptic digest (QC sample)

- Identify scheduling time for target peptides/transitions

  - LC-MRM method preparation

    - Load the Skyline file containing peptides and transitions that will be monitored

    - In Lab Solutions, run the MRM Optimization Wizard to adjust precursor m/z, optimize voltage (select 5 transitions), and adjust product m/z of all heavy and light peptides

    - Run mixture of all heavy and light peptides twice using LC-MRM method created from MRM Optimization Wizard results (set max dwell time to 10 ms)

    - Import .lcd file into Skyline and refine transitions (3 transitions/precursor)

    - Update Lab Solutions .lcm LC-MRM method based on refined transitions from Skyline

  - Timing the peptide detection

    - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

    - Inject the QC sample three times

    - Run mixture of all heavy and light peptides twice

    - Import .lcd file into Skyline and check the automatic integration of all peaks

    - Export the scheduled transition list using 180 sec scheduled windows

    - Update Lab Solutions .lcm LC-MRM method based on scheduled windows determined by Skyline