# STANDARD OPERATING PROCEDURE

Title: Liquid Chromatography, Thermo Scientific EASY-nLC 1200 LC System. Version #: 2; Author: Akhilesh Pandey Lab. Date: 05/01/2016.

# Purpose

The purpose of this document is to describe the liquid chromatography (LC) method for quantitative mass spectrometry-based analyses.

# Scope

This procedure describes the setup of the LC and the chromatographic method parameters. It is specific to the operation of the Thermo Scientific EASY-nLC 1200 LC system.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- HPLC: Thermo Scientific EASY-nLC 1200 LC system (Thermo Fisher Scientific; cat. # LC140)

# Materials

- Injection loop: 20 µL Dionex nanoViper sample loop (Thermo Fisher Scientific; cat. # 6826.2420)

- Autosampler: Autosampler integrated in Easy-nLC 1200 LC system (Thermo Fisher Scientific; cat. # LC140)

- Trap Column: 100 µm I.D. x 2 cm nanoViper column packed with Acclaim PepMap RSLC C18, 5 µm 100 Å (Thermo Fisher Scientific; cat. # 164564)

- Analytical Column: EASY-spray column, 75 µm I.D. x 50 cm nanoViper column packed with Acclaim PepMap RSLC C18, 2 µm 100 Å (Thermo Fisher Scientific; cat. # ES803)

- Autosampler 96-well plate: PCR plate, 150 µL max. well volume (Eppendorf; cat. # 951020401)

- Autosampler 96-well plate cover: Adhesive free film (Excel Scientific; cat. # ZAF-PE-50)

# Reagents

- Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)

- Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)

- Formic Acid: Pierce Formic Acid, LC-MS grade (Pierce; cat. # 28905)

# Solutions

- Mobile phase A: 0.1% formic acid in water

- Mobile phase B: 0.1% formic acid in 95% ACN

# Procedure

- Sample Pickup

  - Draw Speed =                     	20 [µL/s]

  - Wash Volume =                 	100.000 [µL]

  - Rinse Between Reinjections = No

  - Inject Volume = 	       	 10.000 µL

- Pre-column equilibration

  - Max Pressure = 		500 Bar

  - Flow Rate =			unspecified

  - Volume =			25 µL

- Analytical column equilibration

  - Max Pressure = 		750 Bar

  - Flow Rate =			unspecified

  - Volume =			6 µL

- Sample Loading

  - Max Pressure = 		500 Bar

  - Flow gradient =		Isocratic; 100% Mobile phase A

  - Load volume =		25 µL

- Gradient method:

  - Flow rate = 			300 nL/min

  - Temperature = 		45°C

  - Timetable

| Retention (min) | Flow (nL/min) | %B |
| --- | --- | --- |
| 0:00 | 250 | 5 |
| 02:00 | 250 | 7 |
| 52:00 | 250 | 30 |
| 55:00 | 250 | 100 |
| 60:00 | 250 | 100 |

# Referenced Documents

N/A