# STANDARD OPERATING PROCEDURE

Title: Parallel Reaction Monitoring (PRM) Mass Spectrometry, Orbitrap Fusion Lumos. Version #: 1; Author: Akhilesh Pandey Lab. Date: 01/15/2016.

# Purpose

The purpose of this document is to describe the mass spectrometry method for the quantitative analysis of peptides using parallel reaction monitoring (PRM).

# Scope

This procedure describes the setup of the mass spectrometer and the PRM method parameters for the Orbitrap Fusion Lumos. LC parameters are contained in a separate document.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Source: EASY-Spray (Thermo Fisher Scientific)

- Emitter tip: Precision positioned glass emitter integrated into EASY-Spray column (Thermo Fisher Scientific)

- LC-to-source connection: nanoViper fitting; integrated zero dead volume union within EASY-Spray column (Thermo Fisher Scientific)

- Mass spectrometer: Orbitrap Fusion Lumos (Thermo Fisher Scientific)

# Materials

- Injection loop: 20 µL nanoViper sample loop (Thermo Fisher Scientific; cat. # 6826.2420)

- Easy nLC 1200 system (Thermo Fisher Scientific) Nano Flow system with integrated autosampler.

- Trap Column: 100 µm I.D. x 2 cm packed with Acclaim PepMap 100 5 µm, 100 Å C18 (Thermo Fisher Scientific; cat. # 164564)

- Analytical Column: 75 µm I.D. x 50 cm EASY-Spray column packed with Acclaim PepMap RSLC 100 Å  C18, 2 µm (Thermo Fisher Scientific; cat. # ES803)

# Reagents

- Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)

- Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)

- Formic Acid: LC-MS Ultra (Sigma-Aldrich; cat. # 14265)

# Solutions

- Loading pump, mobile phase A: 0.1% formic acid in water

- NanoFlow pump, mobile phase A: 0.1% formic acid in water

- NanoFlow pump, mobile phase B: 0.1% formic acid in 95% ACN

# Procedure

- Setup MS method and tune file parameters

  - NSI Source/Gas parameters

    - Spray Voltage: 2200 V

    - Capillary Temperature: 200 °C

    - Sheath Gas: 0

    - Auxiliary Gas: 0

    - Sweep Gas: 2

  - MS parameters

    - Microscans: 1

    - Resolution: 30,000

    - AGC target: 1e5

    - Isolation window: 2.0 m/z

    - NCE: 32

    - Lock masses: 445.12003 m/z

- Test system suitability with 200 fmol bovine serum albumin tryptic digest once column is conditioned

- Preparing isolation list for target peptides

  - LC-PRM method preparation

    - Load the Skyline file containing peptides that will be monitored during the QC analysis

    - Export the unscheduled transition list as a single method

    - Import the unscheduled transition list as an Inclusion list in a Targeted-MS2  (tMS2) acquisition method in Xcalibur with the parameters set as indicated in Step 1, above

  - Peptide detection

    - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

    - Inject the QC sample 3x

    - Import the data files into the Skyline file

    - Check the automatic integration of all peaks

  - QC sample analysis

    - Import the isolation list into an Inclusion list in a Targeted-MS2  acquisition method in Xcalibur with the parameters set as indicated in Step 1, above

    - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

    - Inject the QC sample 3x

# Referenced Documents

N/A