# STANDARD OPERATING PROCEDURE

Title: MRM mass spectrometry, 5500 QTRAP. Version #: 2; Author: Paulovich lab. Date: 4/10/2013.

# Purpose

The purpose of this document is to describe the mass spectrometry (MS) method for quantitative analysis of peptides using multiple reaction monitoring (MRM).

# Scope

This procedure encompasses the setup of the MS and method parameters. LC parameters are contained in a separate document.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Mass spectrometer: ABSciex 5500 QTRAP

- Source: ADVANCE CaptiveSpray Source for ABSciex (Michrom Bioresources/Bruker)

- Emitter tip: CaptiveSpray tapered tip 20um ID (Michrom Bioresources, SS9/25000/20)

- LC-to-source connection: PEEKsil, 1/32” x 25um x 20cm (Upchurch, 32520)

# Procedure

- Setup MS method parameters

  - Source/Gas Parameters:

    - Curtain Gas (CUR): 10

    - IonSpray Voltage (IS): 1200

    - Ion Source Gas 1 (GS1): 0

    - Ion Source Gas 2 (GS2): 0

    - Interface Heater Temperature (IHT): 110

  - Scheduled MRM Parameters:

    - MRM detection window (sec): 150

    - Target Scan Time (sec): 1.5

  - MS Parameters:

    - Declustering Potential (DP): 100 (or from Skyline)

    - Entrance Potential (EP): 10

    - Collision Energy (CE): From Skyline (default 5500 regression)

    - Collision Cell Exit Potential Q1 (CXP): 10

  - Advanced MS Parameters:

    - Resolution Q1: Unit

    - Resolution Q3: Unit

    - Intensity threshold (total count): 0

    - Settling time (ms): 0

    - Pause between mass ranges (ms): 3

- Test system suitability with appropriate standard once column is conditioned.

- Identify scheduling times for target peptides/transitions

  - Interlab target LC-SRM method preparation

    - Load the Skyline file containing peptides and transitions that will be monitored during the QC analysis.

    - In the Skyline file under Settings/Transition Settings/Predictions, select ‘ABI 5500 QTrap DQ’ under ‘Collision energy:’ and ‘Static’ under ‘Declustering potential’.

    - Export the unscheduled transition list as ‘Multiple methods’ (6), ignoring proteins, with a dwell time of 10ms.

    - Import the unscheduled transition lists into MRM acquisition method on the 5500 with all other parameters set as above (step 1).

  - Timing the peptide detection

    - Set up the autosampler and LC methods as in the accompanying LC SOP.

    - Inject the QC sample 6x.

    - Import the data files into the Skyline file.

    - Check the automatic integration of all peaks.

    - Export the scheduled transition list using above scheduled parameters.

  - QC sample analysis

    - Import the scheduled transition list into an MRM acquisition method on the 5500 with all parameters set as above.

    - Set up the autosampler and LC methods as in the accompanying LC SOP.

    - Inject the QC sample 5 times (or 3x per column for dual column).

  - Instrument performance evaluation

    - Import the data files into the Skyline file.

    - Check the automatic integration of all peaks.

      - Manually adjust integration of peaks, if necessary.

      - Make sure integration start and stop is identical for all transitions of a precursor (go to “Settings”, and check “Integrate All” to enable this feature automatically).

    - Check that peak shape is acceptable.

      - No tailing or fronting.

      - No drop-out of electrospray.

      - No missing transitions.

      - If the peaks are unacceptable, troubleshoot the LC system and re-run the column conditioning procedure.

    - Check that peaks pass reproducibility criteria.

      - Retention time shift is less than 0.5 minutes.

      - Peak area ratio CV is less than 20% for all peaks.

- Analysis of response curve for characterization of assays.

  - LC-MRM method preparation

    - Open Skyline file containing targeted peptides, transitions, and retention times.

    - In the Skyline file under Settings/Transition Settings/Predictions, select ‘ABI 5500 QTrap DQ’ under ‘Collision energy:’ and ‘Static’ under ‘Declustering potential’.

    - Export the scheduled transition list.

    - Import the scheduled transition list into an MRM acquisition method on the 5500 with all parameters set as above (step 1).

    - Set up the autosampler and LC methods as in the accompanying LC SOP.

  - Run order

    - Filenaming convention follows:

      - project_site_assaygroup_date_concentration_replicate#

      - data file example: GO_Hutch_Interlab_ 041112_B2.