# STANDARD OPERATING PROCEDURE

Title: Liquid Chromatography, Waters nanoACQUITY UPLC system. Version #: PRISM; Author: PNNL lab. Date: 07/20/2016.

# Purpose

The purpose of this document is to describe the Liquid Chromatography (LC) method of 1st dimension PRISM fractionation for development of CPTAC PRISM-SRM assays.

# Scope

This procedure is designed to help the setup of LC gradient and method parameters on Waters nanoACQUITY UPLC system of the PRISM fractionation method.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

Milli-Q Integral Water Purification System (EMD Millipore, Darmstadt, Germany, Catalog # ZRXQ010TO)

Collect PAL (LEAP Technologies, Fort Lauderdale, FL, Catalog # MXY013-02A)

nanoACQUITY UPLC (Waters, Milford, MA, Catalog # 176016000)

# Materials

200 μm i.d. × 50 cm long column packed in house with Jupiter® 3 µm C18 300 Å (Phenomenex, Torrance, CA)

twin-tec PCR Plate 96, skirted, colorless  (Eppendorf AG, Hamburg, Germany, Catalog # 951020401)

# Reagents

Acetonitrile, Optima LC/MS Grade (Fisher Scientific, Asheville, NC, Catalog # A955-4)

Acetonitrile/Formic Acid 0.1% Mobile Phase (EMD Millipore, Darmstadt, Germany, Catalog # FX0437P-1)

Ammonium Hydroxide Solution, 28.0-30.0% (Sigma Aldrich, St. Louis, MO, Catalog # 221228)

Pierce™ Formic Acid, LC-MS Grade (Thermo Fisher Scientific, Waltham, MA, Catalog # 28905)

Ammonium Formate, for HPLC >=99.0% (Sigma Aldrich, St. Louis, MO, Catalog # 540-69-2)

# Solutions

Mobile phase A: 10mM Ammonium formate in H2O, pH 9

Mobile phase B: 10mM Ammonium formate in 90% ACN, pH 9

Strong Needle Wash buffer: 100% ACN

Weak Needle Wash buffer: 0.1% FA in H2O

Seal Wash buffer: 10% ACN in H2O

# Procedure

PRISM fractionation (see details in Reference 1)

- nanoACQUITY autosampler parameters:

Under General tab:

‘Partial loop’ under ‘sample loop option’

Wash Solvent: Weak 600 ul, Strong 200 ul

Autosampler temperature: 4 ºC

- PRISM fractionation gradient methods:

- Column Temperature: room temperature

- 45 uL of each sample (0.5 μg/μL) were loaded onto the column and separated using a binary gradient of 1-10%B in 2 min, 10-15% B in 15 min, 15-25% B in 35 min, 25-35% B in 25 min, 35-45% B in 13 min and 45-90% B in 25 min. The flow rate is 2.2 ul/min

- Collection:

- Following the LC separation, the eluate from the capillary column is collected every 1min into a 96 well plate using a Collect Pal system over the course of 96 min.

- Prior to peptide fraction collection, 18 μL of water is added to each well in the plate to avoid peptide loss and also to dilute the peptide fraction for LC-SRM analysis.

- The fraction containing a target peptide is intelligently selected based on the retention time of the peptide obtained by on-line monitoring.

- The fractions are transferred to individual Waters glass vials and stored in autosampler. Now they are ready for 2nd dimension LC-SRM analysis.

- The following is a list of peptides and their locations in terms of fraction number.

| Protein | Peptide | ResponseCurveSamples | ValidationSamples |
| --- | --- | --- | --- |
| LAX1_HUMAN | TDDPGTHVQCVK | 6 | 5 |
| CTNA2_HUMAN | IASSEFADDPCSSVK | 7 | 6 |
| RAF1_HUMAN | LVADCVK | 12 | 10 |
| AMPD1_HUMAN | YNPVGASELR | 12 | 11 or 10 |
| FOSB_HUMAN | IPYEEGPGPGPLAEVR | 16 | 14 |
| SLAF7_HUMAN | VDFPDGGYSLK | 22 | 20 |
| AMPD1_HUMAN | TDNLPENLGYHLK | 23 | 22 |
| CTNA2_HUMAN | WDDSGNDIIVLAK | 24 | 23 |
| SRF_HUMAN | ALIQTCLNSPDSPPR | 28 | 25 |
| DUS1_HUMAN | DGTLALAAGALCR | 37 | 35 |
| ARI4B_HUMAN | LGGFDNIESGAVWK | 40 | 39 |
| ATF3_HUMAN | NLFIQQIK | 44 | 42 |
| ELMO1_HUMAN | LLDLENIQIPDAPPPIPK | 52 | 51 |
| ELMO2_HUMAN | LLDLENIQIPEAPPPIPK | 53 | 51 |

# Referenced Documents

- Shi T, TL Fillmore, X Sun, R Zhao, AA Schepmoes, M Hossain, F Xie, S Wu, JS Kim, NJ Jones, RJ Moore, L Pasa-Tolic, J Kagan, KD Rodland, T Liu, K Tang, DG Camp, II, RD Smith, and W Qian. 2012. "Antibody-free, targeted mass-spectrometric approach for quantification of proteins at low picogram per milliliter levels in human plasma/serum." Proceedings of the National Academy of Sciences of the United States of America 109(38):15395-15400.  doi:10.1073/pnas.1204366109