# STANDARD OPERATING PROCEDURE

Title: MRM Mass Spectrometry, TSQ Vantage. Version #: 1; Author: PNNL Lab. Date: 4/23/2015.

# Purpose

The purpose of this document is to describe the Mass Spectrometry (MS) method for developing peptide multiple reaction monitoring (MRM) assays.

# Scope

This procedure encompasses the setup of the MS and method parameters on Thermo TSQ Vantage.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

Source: in-house built nano-sprary source

Emitter tip: In-house made emitter (20um ID, 360um OD)

Mass spectrometer: TSQ Vantage Triple-Stage Quadrupole Mass Spectrometer (Thermo Fisher Scientific, Waltham, MA)

# Materials

# Reagents

# Procedure

- Setup MS method parameters:

- Source Parameters:

- Ion Spray Voltage: 2400 V

- Capillary Temperature: 325 ºC

- Scheduled MRM Parameters:

- MRM detection window (sec): 300

- Dwell time: 10 ms

- MS parameters:

- Use tuned S-lens value

- Collision gas pressure: 1.5 mTorr

- Q1 and Q3 unit resolution

- Collision energy optimized by LC ramping for each transition

- Intensity threshold: 500

- Test system suitability with appropriate QC standard once column is conditioned.

- Optimize collision energy

- Load the Skyline file containing target peptides and transitions that will be monitored in MRM experiment

- In the Skyline file under Settings/Transition Settings/Predictions, select ‘Thermo TSQ Vantage’ under ‘collision energy’

- Export the transition list with ‘collision energy’ under ‘Optimizing’

- Import the unscheduled transition list into MRM acquisition method on the TSQ with all other parameters set as above (step 1)

- Import TSQ raw file into Skyline and obtain optimal collision energy for each transition

- Identify scheduling times for target peptides/transitions

- Set up the autosampler and LC methods as in the accompanying LC SOP file

- Check the overall status by injecting QC sample

- Export the transition list from Skyline with optimal collision energy

- Inject target peptide mixture into TSQ Vantage

- Import the data files into Skyline and manual check the retention time of each peptide

- Export the scheduled transition list and optimal collision energy

- Method performance evaluation

- Spike the target peptide mixture into real complex matrix (cell lysate, tissue digest, plasma etc)

- Inject the mixed sample

- Import the data file into Skyline

- Check the LC and transition condition:

- Make sure scheduling is ok, no missing peaks

- Make sure peak shape is acceptable, no tailing or fronting

- Make sure each transition is ok, no obvious interference

# Referenced Documents

List any publications or documents referenced in the SOP.