STANDARD OPERATING PROCEDURE

Title: Multiple Reaction Monitoring (MRM) using a TSQ Vantage triple quadrupole mass spectrometer. Version #: 1; Author: Hui Zhang Laboratory – Johns Hopkins University. Date: 04/18/2017.

# Purpose

The purpose of this document is to describe the mass spectrometry method for the quantitative analysis of peptides using multiple reaction monitoring (MRM).

# Scope

This procedure describes the setup of the mass spectrometer and the MRM method parameters for the TSQ Vantage triple quadrupole mass spectrometer. LC parameters are contained in a separate document.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Mass spectrometer: TSQ Vantage (Thermo Scientific)

- Ion source: HESI (Thermo Scientific)

# Procedure

- Setup MS method and tune file parameters

  - ESI Source/Gas parameters

    - Spray voltage: 4000 V

    - Capillary temperature: 270 °C

    - Sheath gas pressure (Arb): 9

    - Collision gas pressure: 1.5 mTorr

  - MS parameters

    - Q1 & Q3 peak width (FWHM): 0.70

    - Cycle time: 0.320 sec

    - Declustering potential: 2 V

    - Chrom filter peak width: 30.0 sec

    - Scheduled window: 120 sec

- Test system suitability with 100 fmol 6 protein digest standard, equimolar (Thermo Scientific; cat. # 88342) – QC sample

- LC-MRM method preparation and execution

  - Set up the autosampler and LC methods as indicated in the Liquid Chromatography SOP

  - Inject the QC sample three times to assess LC-MRM-MS performance

  - Create Skyline file containing .fasta files of assay targets

  - Export unscheduled transition lists (≤ 60 transitions per list)

  - Run heavy target peptides using unscheduled methods and obtain retention times

  - Import .raw files into Skyline and refine transitions

  - Add light peptides into Skyline document

  - Export a scheduled MRM-MS transition list with a 2 min retention time window. (see the Table.1)

  - Run the scheduled method twice

  - Import the raw files into Skyline and check the peak integration

  - Run QC sample at least once every 24 hours

Table.1 The scheduled MRM method

| Precursor m/z | Product m/z | Collision Energy | Start time | End time | Comment |
| --- | --- | --- | --- | --- | --- |
| 779.358 | 838.405 | 26 | 20.52 | 22.52 | NFNASWNSFQNTK |
| 779.358 | 1024.485 | 26 | 20.52 | 22.52 | NFNASWNSFQNTK |
| 779.358 | 1111.517 | 26 | 20.52 | 22.52 | NFNASWNSFQNTK |
| 783.365 | 846.42 | 26 | 20.52 | 22.52 | NFNASWNSFQNTK |
| 783.365 | 1032.499 | 26 | 20.52 | 22.52 | NFNASWNSFQNTK |
| 783.365 | 1119.531 | 26 | 20.52 | 22.52 | NFNASWNSFQNTK |
| 708.888 | 732.389 | 24 | 24.51 | 26.51 | VANLTSVVNSLDGK |
| 708.888 | 918.489 | 24 | 24.51 | 26.51 | VANLTSVVNSLDGK |
| 708.888 | 1019.537 | 24 | 24.51 | 26.51 | VANLTSVVNSLDGK |
| 712.895 | 740.403 | 24 | 24.51 | 26.51 | VANLTSVVNSLDGK |
| 712.895 | 926.503 | 24 | 24.51 | 26.51 | VANLTSVVNSLDGK |
| 712.895 | 1027.551 | 24 | 24.51 | 26.51 | VANLTSVVNSLDGK |
| 928.481 | 774.41 | 31 | 26.08 | 28.08 | GNLTYGYVTILNGSDIR |
| 928.481 | 988.542 | 31 | 26.08 | 28.08 | GNLTYGYVTILNGSDIR |
| 928.481 | 1087.611 | 31 | 26.08 | 28.08 | GNLTYGYVTILNGSDIR |
| 933.485 | 784.419 | 31 | 26.08 | 28.08 | GNLTYGYVTILNGSDIR |
| 933.485 | 998.55 | 31 | 26.08 | 28.08 | GNLTYGYVTILNGSDIR |
| 933.485 | 1097.619 | 31 | 26.08 | 28.08 | GNLTYGYVTILNGSDIR |
| 928.973 | 774.41 | 31 | 26.16 | 28.16 | GN[+1.0]LTYGYVTILNGSDIR |
| 928.973 | 988.542 | 31 | 26.16 | 28.16 | GN[+1.0]LTYGYVTILNGSDIR |
| 928.973 | 1087.611 | 31 | 26.16 | 28.16 | GN[+1.0]LTYGYVTILNGSDIR |
| 933.977 | 784.419 | 31 | 26.16 | 28.16 | GN[+1.0]LTYGYVTILNGSDIR |
| 933.977 | 998.55 | 31 | 26.16 | 28.16 | GN[+1.0]LTYGYVTILNGSDIR |
| 933.977 | 1097.619 | 31 | 26.16 | 28.16 | GN[+1.0]LTYGYVTILNGSDIR |
| 797.419 | 486.303 | 27 | 27.52 | 29.52 | M[+16.0]IENGSLSFLPTLR |
| 797.419 | 599.388 | 27 | 27.52 | 29.52 | M[+16.0]IENGSLSFLPTLR |
| 797.419 | 833.488 | 27 | 27.52 | 29.52 | M[+16.0]IENGSLSFLPTLR |
| 802.423 | 496.312 | 27 | 27.52 | 29.52 | M[+16.0]IENGSLSFLPTLR |
| 802.423 | 609.396 | 27 | 27.52 | 29.52 | M[+16.0]IENGSLSFLPTLR |
| 802.423 | 843.496 | 27 | 27.52 | 29.52 | M[+16.0]IENGSLSFLPTLR |