| CPTAC STANDARD OPERATING PROCEDURE |  |  |
| --- | --- | --- |
| Title: Liquid Chromatography Mass Spectrometry Using Parallel Reaction Monitoring for CPTAC Experiments 1 and 2 |  |  |
| SOP: WU-SOP-MS-PRM |  |  |
| Version1 2 | Author: Townsend Lab |  |
| Date: March, 2021 |  |  |

# Purpose

The purpose of this document is to describe the mass spectrometric methods for quantitative analysis of standard synthetic peptide admixtures using stable isotope dilution mass spectromeryas described for CPTAC Experiments 1 and 2 (https://assays.cancer.gov/guidance-document/).

Scope

This procedure encompasses the optimization of a nano-LC-MS method for the scheduled acquisition of full scan MS2 spectra for pairs of high purity natural abundance and stable isotope labeled peptides. The method (PRM, parallel reaction monitoring) to acquire the data for reverse  and forward response curves as described in WU-SOP-EXP1-2021 using a Q-Exactive™ hybrid quadrupole mass spectrometer. The method optimization and analysisof the LC-MS data  were performed with Skyline software (https://brendanx- uw1.gs.washington.edu/labkey/project/home/software/Skyline/begin.view). The procedures for calibrating and determining the performance of the system using standard benchmark peptides and a standard complex peptide mixture are described in WU-SOP-SETUP-2021. The gradient methods for both Experiments 1 and 2 are described in WU-SOP-LC-1-2021.

Page 1 of 6

Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures and relevant good practice procedures of laboratory biochemistry and analytical chemistry.  The interpretation of results must be performed by person(s) trained in the procedure and trained in the interpretation of these data.

# Materials

# Reagents

- Acetonitrile (MeCN) (Honeywell Fluka, cat. no. 34967)

- Formic Acid (FA) (Honeywell Fluka, cat. no. 56302-50ML-F)

- Pierce Retention Time Calibration Mixture (PRTC) (ThermoFisher Scientific, cat. no. 88321)

- Pierce HeLa Protein Digest Standard, (ThermoFisher Scientific, cat. no.  88328)

- Water (Fluka, cat. no. 39253-1L-R).

Equipment

- Q-Exactive quadrupole-Orbitrap™ mass spectrometer (ThermoFisher Scientific, cat. no. IQLAAEGAAPFALGMAZR)

- EASY-Spray™ source (ThermoFisher Scientific, cat. no. ES081)

- EASY-nLC™ 1000 (ThermoFisher Scientific, cat. no. LC120)

- LC-to-source connection: nano Viper™ tubing (ThermoFisher Scientific, cat. no. 6041.5261)

- EASY-Spray Column: 75 µm x 50 cm PepMap™ RSLC C18, 2 µm, 100 Å (ThermoFisher Scientific, cat. no. ES803)

- Injection loop: 20 µL PEEKsil™, 100 µm (ThermoFisher Scientific, cat. no. LC472)

- Micro-centrifuge (Eppendorf, cat. no.  5424R)

- Pipet-lite™ (Rainin, cat. no. XLS, P20, P200, P1000)

- Autosampler vials (Sun-Sri, cat. no. 200 046)

- Autosampler vial caps (Sun-Sri, cat. no. 501 382)

- Clear vials, 4 mL (National Scientific, cat. no. B7990-2)

- Axygen® MAXYmum™ P200 recovery tips (Axygen, cat. no. T-200-C-L-STK)

- Axygen® MAXYmum™ P20 recovery tips (Axygen, cat. no. T-20-C-L-STK)

- Axygen® MAXYmum™ P1000 recovery tips (Axygen, cat. no. T-1000-C-L-R)

# Software

- Skyline (MacCross Lab Software, https://skyline.ms/project/home/software/Skyline/begin.view)

## Page 2 of 6

# Procedure

Identifying scheduling times for targeted peptides

(NOTE: The assessment of LC-MS system performance is documented in WU-SOP-MS-SETUP-2021)

- Place vial containing equimolar H and L solution of all standard peptides into one autosampler (solution preparation in WU-SOP-EXP1-2021).

- Begin data acquisition with optimized instrument according to WU-SOP- MS-SETUP-2021.

Building a scheduled method

- Set up the autosampler and LC methods as described in WU-SOP-LC-2021 for either Experiment 1 (WU-SOP-EXP1-2021) or Experiment 2 (WU-SOP-EXP2-2021).

- Inject 2.5 µl of the QC sample containing equimolar H/L standard peptide admixtures of L/H standard peptides in sample matrix and acquire the DDA data. The preparation of the equimolar sample is described in WU-SOP-EXP2-2021.

- Convert and search file and build spectral library of the .dat files in Skyline document.

- Load raw data file into skyline document.

- Manually check the automatic peak selection/integration of all peaks. Adjust if necessary. See Skyline tutorial, “Targeted Method Refinement” for reference. https://brendanx-uw1.gs.washington.edu/labkey/wiki/home/software/Skyline/page. view?name=tutorial_method_refine).

- Export an isolation list for the PRM targets under File/Export/Isolation List. Set instrument as Q-Exactive, single method and for Method Type, select Scheduled. See Skyline tutorial, “Custom & Live Reports” for reference. https://brendanx- uw1.gs.washington.edu/labkey/wiki/home/software/Skyline/page. view?name=tutorial_custom_reports.

Scheduled method testing

- Import the scheduled precursor m/z list as an inclusion list under global settings in the DIA method on the QExactive instrument with parameters set as above (steps 1 and 2).

- Set up the autosampler and LC methods as described in WU-SOP- LC-2021.

- Inject a (2.5 µl) ‘HIGH’ quality control sample containing equimolar H/L standard peptides (WU-SOP-EXP2-2021) spiked in assay matrix and acquire the scheduled PRM-MS data.

Instrument performance evaluation with scheduled method

- Manually adjust integration of peaks, if necessary. See Skyline tutorial, “Targeted Method Refinement” for reference. https://brendanx- uw1.gs.washington.edu/labkey/wiki/home/software/Skylin e/page.view?name=tutorial_method_refine

## Page 3 of 6

- Import the .raw files into the Skyline document.

- Check the automatic peak selection/integration of all peaks.

- Check integration by determining if start and stop are identical for all transitions of a precursor (go to “Settings”, and check box “Integrate All” to enable this feature automatically.

Assess peak shape

- Check peak shape to eliminate tailing or fronting.

- Determine that there is no missing data (e.g. drop-out of electrospray).

- If the peaks are unacceptable, troubleshoot the LC system using the benchmark procedure described above and in WU-SOP-LC-2021 and re-run the column conditioning procedure.

- Apply criteria for peak consistency. Retention time shift should be less than 0.4 min between injections, and peak intensities between injections should show less than 20 % variability.

Final method configuration

- Save Skyline method test file as a final method.

- Remove unscheduled DDAi (Step 5A) and method building scheduled DDAi

raw files (Step 5B) from the document (Edit>Manage Results). Save the document. The resulting Skyline document should now contain the most recent scheduled method testing data files with current retention times.

- Export the scheduled precursor m/z list with retention times as an isolation list using above parameters.

- Import the isolation list into PRM method on the Q-Exactive™ instrument with all other parameters set as above (steps 1 and 2).

Analysis of response curve for characterization of assays

- LC-PRM-MS method preparation: Load the Skyline file containing peptide sequences, precursors, transitions, and retention times to be monitored during the LC-PRM-MS experiment.

- Export the scheduled precursor m/z list as an isolation list (File-Export- Isolation list) and select data files for retention time setting. Set the retention time window to 3 min. See Skyline tutorial, “Custom & Live Reports” for reference. https://brendanx- uw1.gs.washington.edu/labkey/wiki/home/software/Skyline/page.vie w?name=tutorial_custom_reports

- Import the scheduled precursor m/z list as an inclusion list in the Xcalibur Mass spectrometer method on the Q-Exactive™ instrument with all other instrument parameters set as above (steps 1 and 2).

- Set up the autosampler and LC methods as described in WU SOP-LC3.

- Experiment 1 or Experiment 2.

- Refer to WU-SOP-EXP1-2021 or WU-SOP-EXP2-2021 for details on samples and queues.

## Page 4 of 6

| Table 1. PRM Method Parameters |  |
| --- | --- |
| Source/Gas Parameters: | Value |
| IonSpray Voltage (IS) | 1800 |
| S-lens | 60 |
| Interface Heater Temperature | 275°C |
| Scheduled PRM Parameters |  |
| PRM detection window (sec) | 180 |
| Target Scan Time (sec) |  |
| Full MS-SIM Parameters |  |
| Resolution | 70,000 |
| AGC target | 3 x 106 |
| Maximum IT | 60ms |
| Spectrum data type | profile |
| MS2 Parameters (DIA) |  |
| Collision Energy | 27 |
| Isolation window | 2.0 m/z |
| Resolution MS2 | 17,500 |
| AGC target | 2 x 105 |
| Maximum IT | 60 ms |
| Spectrum data type | profile |
| Loop count | 30 |

Referenced Documents

WU-SOP-MS-SETUP-2021

WU-SOP-LC-2021

WU-SOP-EXP1-2021

WU-SOP-EXP2-2021

Skyline manual and tutorials. https://brendanx- uw1.gs.washington.edu/labkey/wiki/home/software/Skyline/page.view?name=tut orials

Peterson AC, Russell JD, Bailey DJ, Westphall MS, Coon JJ. Parallel reaction monitoring for high resolution and high mass accuracy quantitative, targeted proteomics. Molecular & cellular proteomics: MCP. 2012;11(11):1475-88. Epub 2012/08/07. doi: 10.1074/mcp.O112.020131. PubMed PMID: 22865924; PMCID: Pmc3494192.

ThermoFisher Scientific customer manuals website

Page 5 of 6

# Abbreviations

MeCN, acetonitrile

FA, formic acid

LC-MS, nano-LC interfaced to a high-resolution quadrupole-orbitrap mass spectrometer as described in WU-SOP-LC-1 and WU-SOP-MS-1

H (heavy), stable isotope labeled synthetic peptide

L (light) natural isotope abundance synthetic peptide

PRM, parallel reaction monitoring mass spectrometry

IT, fill time

DDA, data-dependent LC_MS acquisition

DIA, data independent acquisition

IS , Ion Spray

## Page 6 of 6