# Title: Multiplexed mass spectrometry-based assays for profiling cellular growth signaling dynamics

# Protocols for the CPTAC Assay Portal

Standard Procedures for the following methods:

1.	Peptide Selection	1

2.	Materials	1

3.	Cell culture	10

4.	Antibody beads crosslinking	12

5.	Peptide Mastermix Preparation	13

6.	Generation of cell lysates	21

7.	Protein Digestion	23

8.	Immunoaffinity Enrichment	26

9.	Analysis of Samples by LC-MRM	31

10.	Data Analysis	33

11.	Run Orders	34

12.	Fit-for-purpose validation	36

# Peptide Selection

## 1.1 Selection criteria

Using the existing cancer biology literature and proteomics resources, each of these targets was comprehensively reviewed. First, existing LC/MS proteomic datasets (including data from breast, ovarian, and colorectal cancer tissues, cancer cell lines, and public databases) were mined for empirical evidence of detection of corresponding peptides. For modified targets, datasets incorporating phosphopeptide enrichment were used for searching. Each dataset was searched for matching peptides to the gene product identified on the master list, or in the case of post-translationally modified targets, the datasets were searched for the tryptic sequence containing the modification (searches allowed for identifications encompassing missed trypsin cleavage sites). Selected peptides were required to be proteotypic (i.e. unique to the protein target of interest and featuring a good response by mass spectrometry) and were prioritized by frequency of observation, MS intensity, length (between ~7-25aa), hydrophobicity (10-40 by SSRCalc) or retention time, charge state (z=2,3) (45), amino acid composition (deprioritize M, N-terminal Q, N-terminal C, multiple P, previous and next amino acids containing trypsin sites (ie. ragged ends)), and frequency of missed cleavage products. After ranking of the peptides based on those criteria, further review focused on a peptide-level knowledge of biology, which included known sites of mutation using CBioPortal (48), post-translational modifications using PhosphoSitePlus (49, 50), and assessment of potential interference using the in silico Peptide Interference Predictor (51).

# Materials

## 2.1 Materials

Procedures below reference some of the following custom-prepared materials.

| Name | Description | Notes |
| --- | --- | --- |
| Antibody-beads Multimix #1 | mAbs to phosphorylated peptides crosslinked to Protein G beads |  |
| Antibody-beads Multimix #2 | mAbs to unmodified peptides crosslinked to Protein G beads |  |
| Reference Lysate | Pool of undigested cell lysate to be used as quality control and calibrator |  |
| MM1 | Unmodified peptides mastermix, light | 8.011 uM |
| MM2 | Unmodified peptides mastermix, SIS (heavy) | 8.011 uM |
| MM3 | Phosphorylated peptides mastermix, light | 18.856 uM |
| MM4 | Phosphorylated peptides mastermix, heavy | 18.856 uM |
| MM1* | Make-up, high concentration unmod light | 2-10x, varies by pep |
| MM2* | Make-up, high concentration unmod heavy | 2-10x |
| MM3* | Make-up, high concentration phos light | 2-25x |
| MM4* | Make-up, high concentration phos heavy | 2-25x |

All samples for CPTAC fit-for-purpose validation experiments consisted of a background lysate digest mix of the following:

| Cell Lysate | Relative contribution to the pooled lysate digest |
| --- | --- |
| MCF10A-EV | 5 |
| T47D | 1.25 |
| CCRF-CEM | 1 |
| COLO205 | 2 |
| COR-L23 | 2.5 |
| H2444 | 3 |
| H2122 | 3 |
| H1792 | 1.25 |
| HEPG2 | 2 |
| K-562 | 1.25 |
| NCI H226 | 3.75 |

The reference lysate consists of the following pool of lysates (undigested):

| Cell Lysate | Relative contribution to the pooled lysate |
| --- | --- |
| A375 | 1 |
| HCT116 | 1 |
| HT29 | 2 |
| MCF10A- KRAS | 3 |
| SK-Mel-2 | 2 |
| BxPC-3 | 1 |

Spike levels for the response curve samples - reverse response curves: variable heavy, constant light (CPTAC Experiment #1):

|  |  |  |  | Phospho light spike (fmol) (MM4) | Non-modified light spike (fmol) (MM2) |
| --- | --- | --- | --- | --- | --- |
|  |  |  |  | 50 | 20 |
| Curve | Heavy Spike amount (fmol) | Dilution Factor | Concentration (fmol/mg) | Expected Ratio if no endogenous | Expected Ratio if no endogenous |
| I | 2000 |  | 4000 | 40 | 100 |
| H | 200 | 10 | 400 | 4 | 10 |
| G | 20 | 10 | 40 | 0.4 | 1 |
| F | 8 | 2.5 | 16 | 0.16 | 0.4 |
| E | 3.2 | 2.5 | 6.4 | 0.064 | 0.16 |
| D | 1.28 | 2.5 | 2.56 | 0.0256 | 0.064 |
| C | 0.512 | 2.5 | 1.024 | 0.01024 | 0.0256 |
| B | 0.2048 | 2.5 | 0.4096 | 0.004096 | 0.01024 |

Spike levels for the QC validation samples (CPTAC Experiment #2)

| QC sample | NP heavy | P heavy | NP Expected ratio | P Expected ratio |
| --- | --- | --- | --- | --- |
| Low | 200 | 500 | 10 | 10 |
| Med | 20 | 50 | 1 | 1 |
| High | 2 | 5 | 0.1 | 0.1 |

Phosphopeptide table: phosphopeptides included in mastermixes are indicated for mastermix 3 (lights) and mastermix 4 (heavy), phosphopeptides present in makeup mixes (MM3* and MM4*) are indicated by an “X”, the antibodies included in Ab multimix 1 are indicated by Antibody ID, and peptides to monitor are indicated by a “1” in the MRM method # column. For the phosphopeptides MRM method, be sure to include the unmodified targets indicated in the table for the Unmodified peptides (below).

| Protein | Peptide | Modification site | Light Peptide Mastermix | Heavy Peptide Mastermix | Present in Makeup Peptide Mix | MRM met-hod # |
| --- | --- | --- | --- | --- | --- | --- |
| AKT1 | (pT)FCGTPEYLAPEVLEDNDYGR | pT308 | MM3 | MM4 | X | 1 |
| AKT1 | RPHFPQF(pS)YSASGTA | pS473 | MM3 | MM4 |  | 1 |
| AKT1S1 | LN(pT)SDFQK | pT246 | MM3 | MM4 |  |  |
| AKT2 | THFPQF(pS)YSASIR | pS473 | MM3 | MM4 |  | 1 |
| AKT2 | THFPQF(pS)YSASIRE | pS473 | MM3 | MM4 | X | 1 |
| AKT3 | RPHFPQF(pS)YSASGR | pS472 | MM3 | MM4 |  | 1 |
| AKT3 | RPHFPQF(pS)YSASGRE | pS472 | MM3 | MM4 |  | 1 |
| ARAF | D(pS)GYYWEVPPSEVQLLK | pS299 | MM3 | MM4 | X | 1 |
| ARAF | NLGYRD(pS)GYYWEVPPSEVQLLK | pS299 | MM3 | MM4 | X | 1 |
| BRAF | D(pS)SDDWEIPDGQITVGQR | pS446 | MM3 | MM4 |  | 1 |
| BRAF | GDGGSTTGLSA(pT)PPASLPGSLTNVK | pT401 | MM3 | MM4 |  | 1 |
| BRAF | RD(pS)SDDWEIPDGQITVGQR | pS446 | MM3 | MM4 |  | 1 |
| CTNNB1 | RT(pS)MGGTQQQFVEGVR | pS552 | MM3 | MM4 |  |  |
| CTNNB1 | T(pS)MGGTQQQFVEGVR | pS552 | MM3 | MM4 |  |  |
| EGFR | GSHQISLDNPD(pY)QQDFFPK | pY1172 | MM3 | MM4 | X | 1 |
| EGFR | GSTAENAE(pY)LR | pY1197 | MM3 | MM4 |  | 1 |
| EGFR | (pY)SSDPTGALTEDSIDDTFLPVPEYINQSVPK | pY1069 | MM3 | MM4 | X | 1 |
| EGFR | Y(pS)(pS)DPTGALTEDSIDDTFLPVPEYINQSVPK | S1070/S1071 | MM3 |  | LIGHT ONLY (MM3*) |  |
| EGFR | YSSDPTGALTEDSIDDTFLPVPE(pY)INQSVPK | pY1092 | MM3 | MM4 | LIGHT ONLY (MM3*) | 1 |
| EIF2A | (pS)PDLAPTPAPQSTPR | pS505 | MM3 | MM4 |  |  |
| EIF2A | SPDLAPTPAPQS(pT)PR | pT518 | MM3 | MM4 |  |  |
| EIF2A | SDK(pS)PDLAPTPAPQSTPR | pS505 | MM3 | MM4 |  |  |
| EIF2A | SDKSPDLAPTPAPQS(pT)PR | pT518 | MM3 | MM4 |  |  |
| ERBB2 | GTPTAENPE(pY)LGLDVPV | pY1248 | MM3 | MM4 | X | 1 |
| ERBB2 | LLDIDETE(pY)HADGGK | pY877 | MM3 | MM4 |  | 1 |
| FOXO1 | AA(pS)MDNNSK | pS256 | MM3 | MM4 | X | 1 |
| FOXO3 | AV(pS)MDNSNK | pS253 | MM3 | MM4 | X | 1 |
| FOXO3 | AV(pS)MDNSNKYTK | pS253 | MM3 | MM4 |  | 1 |
| GSK3B | TT(pS)FAESCKPVQQPSAFGSMK | pS9 | MM3 | MM4 |  | 1 |
| MAP2K2 | L(camC)DFGVSGQLID(pS)MANSFVGTR | pS221 | MM3 | MM4 | X |  |
| MAPK1 | VADPDHDHTGFL(pT)E(pY)VATR | pT185/pY187 | MM3 | MM4 |  | 1 |
| MAPK1 | VADPDHDHTGFL(pT)EYVATR | pT185 | MM3 | MM4 |  | 1 |
| MAPK1 | VADPDHDHTGFLTE(pY)VATR | pY187 | MM3 | MM4 |  | 1 |
| MAPK14 | HTDDEM(pT)G(pY)VATR | pT180/pY182 | MM3 | MM4 |  |  |
| MAPK14 | HTDDEM(pT)GYVATR | pT180 | MM3 | MM4 |  |  |
| MAPK14 | HTDDEMTG(pY)VATR | pY182 | MM3 | MM4 |  |  |
| MAPK3 | IADPEHDHTGFL(pT)E(pY)VATR | pT202/pY204 | MM3 | MM4 |  | 1 |
| MAPK3 | IADPEHDHTGFL(pT)EYVATR | pT202 | MM3 | MM4 |  | 1 |
| MAPK3 | IADPEHDHTGFLTE(pY)VATR | pY204 | MM3 | MM4 |  | 1 |
| MAPK8 | TAGTSFMM(pT)P(pY)VVTR | pT183/pY185 | MM3 | MM4 |  | 1 |
| MAPK8 | TAGTSFMM(pT)PYVVTR | pT183 | MM3 | MM4 |  | 1 |
| MAPK8 | TAGTSFMMTP(pY)VVTR | pY185 | MM3 | MM4 |  | 1 |
| MTOR | TGTTVPESIH(pS)FIGDGLVKPEALNK | pS2481 | MM3 | MM4 |  | 1 |
| MTOR | TGTTVPESIH(pS)FIGDGLVKPEALNKK | pS2481 | MM3 | MM4 | X | 1 |
| MTOR | TD(pS)YSAGQSVEILDGVELGEPAHK | pS2448 | MM3 | MM4 | X | 1 |
| MTOR | TRTD(pS)YSAGQSVEILDGVELGEPAHK | pS2448 | MM3 | MM4 | X | 1 |
| RAF1 | ST(pS)TPNVHMVSTTLPVDSR | pS259 | MM3 | MM4 | X | 1 |
| RPS6KA1 | GF(pS)FVATGLMEDDGKPR | pS380 | MM3 | MM4 |  |  |
| STAT3 | FI(camC)VTPTT(camC)SNTIDLPM(pS)PR | pS727 | MM3 | MM4 |  |  |
| STAT3 | Y(camC)RPESQEHPEADPGSAAP(pY)LK | pY705 | MM3 | MM4 |  |  |

Unmodified peptide table: peptides included in mastermixes are indicated for mastermix 1 (lights) and mastermix 2 (heavy), peptides present in makeup mixes (MM1* and MM2*) are indicated by an “X”, the antibodies included in Ab multimix 2 are indicated by Antibody ID, and peptides to monitor are indicated by a “1” for MRM method 1 and a “2” for MRM method 2 (in the MRM method # column). For the phosphopeptides MRM method, be sure to include the unmodified targets indicated in the table for the Unmodified peptides (below).

| Protein | Peptide Sequence | Light Peptide Mastermix | Heavy Peptide Mastermix | In Makeup Peptide Mix? | MRM met-hod# |
| --- | --- | --- | --- | --- | --- |
| AKT1 | FFAGIVWQHVYEK | MM1 | MM2 | X | 2 |
| AKT1 | FYGAEIVSALDYLHSEK | MM1 | MM2 |  |  |
| AKT1 | RPHFPQFSYSASGTA | MM1 | MM2 |  | 1 |
| AKT1 | SLLSGLLK | MM1 | MM2 |  | 2 |
| AKT1 | TFCGTPEYLAPEVLEDNDYGR | MM1 | MM2 |  | 1 |
| AKT1S1 | LNTSDFQK | MM1 | MM2 |  |  |
| AKT2 | EGISDGATMK | MM1 | MM2 |  |  |
| AKT2 | FYGAEIVSALEYLHSR | MM1 | MM2 |  |  |
| AKT2 | LLPPFKPQVTSEVDTR | MM1 | MM2 | X | 2 |
| AKT2 | SLLAGLLK | MM1 | MM2 |  |  |
| AKT2 | THFPQFSYSASIR | MM1 | MM2 |  | 1 |
| AKT2 | THFPQFSYSASIRE | MM1 | MM2 |  | 1 |
| AKT3 | DEVAHTLTESR | MM1 | MM2 |  | 2 |
| AKT3 | EGITDAATMK | MM1 | MM2 |  | 2 |
| AKT3 | FYGAEIVSALDYLHSGK | MM1 | MM2 |  |  |
| AKT3 | RPHFPQFSYSASGR | MM1 | MM2 |  | 1 |
| AKT3 | RPHFPQFSYSASGRE | MM1 | MM2 |  | 1 |
| AKT3 | TDGSFIGYK | MM1 | MM2 |  |  |
| ARAF | DSGYYWEVPPSEVQLLK | MM1 | MM2 | X | 1 |
| ARAF | GLNQDCCVVYR | MM1 | MM2 |  | 2 |
| ARAF | IGTGSFGTVFR | MM1 | MM2 |  | 2 |
| ARAF | NLGYRDSGYYWEVPPSEVQLLK | MM1 | MM2 | X | 1 |
| ARAF | TQADELPACLLSAAR | MM1 | MM2 |  | 2 |
| ARAF | VSQPTAEQAQAFK | MM1 | MM2 |  |  |
| BRAF | AGFQTEDFSLYACASPK | MM1 | MM2 |  |  |
| BRAF | DSSDDWEIPDGQITVGQR | MM1 | MM2 |  | 1 |
| BRAF | GDGGSTTGLSATPPASLPGSLTNVK | MM1 | MM2 |  | 1 |
| BRAF | GLIPECCAVYR | MM1 | MM2 |  | 2 |
| BRAF | RDSSDDWEIPDGQITVGQR | MM1 | MM2 |  | 1 |
| BRAF | SNNIFLHEDLTVK | MM1 | MM2 |  |  |
| CCND1 | ACQEQIEALLESSLR | MM1 | MM2 | X | 2 |
| CCND1 | AEETCAPSVSYFK | MM1 | MM2 |  | 2 |
| CCND1 | AYPDANLLNDR | MM1 | MM2 |  | 2 |
| CCND1 | FLSLEPVKK | MM1 | MM2 | X | 2 |
| CCND1 | VIKCDPDCLR | MM1 | MM2 |  | 2 |
| CDH1 | GLDARPEVTR | MM1 | MM2 |  | 2 |
| CDH1 | GQVPENEANVVITTLK | MM1 | MM2 |  | 2 |
| CDH1 | NDVAPTLMSVPR | MM1 | MM2 |  |  |
| CDH1 | NTGVISVVTTGLDR | MM1 | MM2 |  |  |
| CDH1 | TAYFSLDTR | MM1 | MM2 |  |  |
| CDH2 | GPFPQELVR | MM1 | MM2 |  |  |
| CDH2 | IDPVNGQITTIAVLDR | MM1 | MM2 |  |  |
| CDH2 | LNGDFAQLNLK | MM1 | MM2 |  |  |
| CDH2 | LSDPANWLK | MM1 | MM2 |  | 2 |
| CDH2 | SAAPHPGDIGDFINEGLK | MM1 | MM2 | X | 2 |
| CTNNB1 | RTSMGGTQQQFVEGVR | MM1 | MM2 |  |  |
| CTNNB1 | TSMGGTQQQFVEGVR | MM1 | MM2 |  |  |
| EGFR | GSHQISLDNPDYQQDFFPK | MM1 | MM2 | X | 1 |
| EGFR | GSTAENAEYLR | MM1 | MM2 |  | 1 |
| EGFR | IPLENLQIIR | MM1 | MM2 |  | 2 |
| EGFR | NLQEILHGAVR | MM1 | MM2 |  | 2 |
| EGFR | YSSDPTGALTEDSIDDTFLPVPEYINQSVPK | MM1 | MM2 | X | 1 |
| EIF2A | SDKSPDLAPTPAPQSTPR | MM1 | MM2 |  |  |
| EIF2A | SPDLAPTPAPQSTPR | MM1 | MM2 |  |  |
| ERBB2 | ELVSEFSR | MM1 | MM2 |  |  |
| ERBB2 | GIWIPDGENVK | MM1 | MM2 |  | 2 |
| ERBB2 | GTPTAENPEYLGLDVPV | MM1 | MM2 | X | 1 |
| ERBB2 | LLDIDETEYHADGGK | MM1 | MM2 |  | 1 |
| ERBB3 | ANDALQVLGLLFSLAR | MM1 | MM2 |  |  |
| ERBB3 | GDSAYHSQR | MM1 | MM2 |  |  |
| ERBB3 | LTFQLEPNPHTK | MM1 | MM2 |  | 2 |
| ERBB3 | SLEATDSAFDNPDYWHSR | MM1 | MM2 |  | 2 |
| ERBB3 | YLERGESIEPLDPSEK | MM1 | MM2 |  | 2 |
| FOS | APHPFGVPAPSAGAYSR | MM1 | MM2 |  | 2 |
| FOS | GSSSNEPSSDSLSSPTLLAL | MM1 | MM2 | X | 2 |
| FOS | LEFILAAHR | MM1 | MM2 |  |  |
| FOS | SALQTEIANLLK | MM1 | MM2 |  |  |
| FOS | TEPFDDFLFPASSR | MM1 | MM2 |  | 2 |
| FOXO3 | AVSMDNSNK | MM1 | MM2 |  |  |
| FOXO3 | AVSMDNSNKYTK | MM1 | MM2 |  |  |
| GSK3B | LLEYTPTAR | MM1 | MM2 |  | 2 |
| GSK3B | QTLPVIYVK | MM1 | MM2 |  |  |
| GSK3B | TPPEAIALCSR | MM1 | MM2 |  | 2 |
| GSK3B | TTSFAESCKPVQQPSAFGSMK | MM1 | MM2 |  | 1 |
| GSK3B | VIGNGSFGVVYQAK | MM1 | MM2 |  | 2 |
| MAP2K1 | IPEQILGK | MM1 | MM2 |  |  |
| MAP2K1 | ISELGAGNGGVVFK | MM1 | MM2 |  | 2 |
| MAP2K1 | VSHKPSGLVMAR | MM1 | MM2 |  |  |
| MAP2K2 | L(camC)DFGVSGQLIDSMANSFVGTR | MM1 | MM2 |  |  |
| MAPK1 | VADPDHDHTGFLTEYVATR | MM1 | MM2 |  | 1 |
| MAPK14 | HTDDEMTGYVATR | MM1 | MM2 |  |  |
| MAPK3 | ALDLLDR | MM1 | MM2 |  | 2 |
| MAPK3 | GQPFDVGPR | MM1 | MM2 |  |  |
| MAPK3 | IADPEHDHTGFLTEYVATR | MM1 | MM2 |  | 1 |
| MAPK3 | NYLQSLPSK | MM1 | MM2 |  |  |
| MAPK3 | SQQLSNDHICYFLYQILR | MM1 | MM2 |  |  |
| MAPK8 | TAGTSFMMTPYVVTR | MM1 | MM2 |  | 1 |
| MTOR | DLELAVPGTYDPNQPIIR | MM1 | MM2 |  | 2 |
| MTOR | IQSIAPSLQVITSK | MM1 | MM2 |  | 2 |
| MTOR | LFDAPEAPLPSR | MM1 | MM2 |  | 2 |
| MTOR | LTESLDFTDYASR | MM1 | MM2 |  | 2 |
| MTOR | TDSYSAGQSVEILDGVELGEPAHK | MM1 | MM2 |  | 1 |
| MTOR | TGTTVPESIHSFIGDGLVKPEALNK | MM1 | MM2 |  | 1 |
| MTOR | TGTTVPESIHSFIGDGLVKPEALNKK | MM1 | MM2 | X | 1 |
| MTOR | TRTDSYSAGQSVEILDGVELGEPAHK | MM1 | MM2 |  | 1 |
| MTOR | VLGLLGALDPYK | MM1 | MM2 |  | 2 |
| PTEN | AQEALDFYGEVR | MM1 | MM2 |  | 2 |
| PTEN | GVTIPSQR | MM1 | MM2 |  | 2 |
| PTEN | IYNLCAER | MM1 | MM2 |  | 2 |
| PTEN | IYSSNSGPTR | MM1 | MM2 |  |  |
| PTEN | NHLDYRPVALLFHK | MM1 | MM2 |  | 2 |
| RAF1 | DSSYYWEIEASEVMLSTR | MM1 | MM2 |  |  |
| RAF1 | GLQPECCAVFR | MM1 | MM2 |  | 2 |
| RAF1 | GYASPDLSK | MM1 | MM2 |  | 2 |
| RAF1 | STSTPNVHMVSTTLPVDSR | MM1 | MM2 |  | 1 |
| RAF1 | VVDPTPEQFQAFR | MM1 | MM2 |  | 2 |
| RASGRF1 | LLYGEPPKSPR | MM1 | MM2 |  |  |
| RASGRF1 | NSLDYAK | MM1 | MM2 |  |  |
| RASGRF1 | SLELLFASGQNNK | MM1 | MM2 |  |  |
| RPS6KA1 | GFSFVATGLMEDDGKPR | MM1 | MM2 |  |  |
| RPTOR | ALETIGANLQK | MM1 | MM2 |  | 2 |
| RPTOR | SLIVAGLGDGSIR | MM1 | MM2 |  |  |
| RPTOR | SYNCTPVSSPR | MM1 | MM2 |  | 2 |
| RPTOR | VLDTSSLTQSAPASPTNK | MM1 | MM2 |  |  |
| RPTOR | VLNSIAYK | MM1 | MM2 |  |  |
| STAT3 | FI(camC)VTPTT(camC)SNTIDLPMSPR | MM1 | MM2 |  |  |
| STAT3 | Y(camC)RPESQEHPEADPGSAAPYLK | MM1 | MM2 |  |  |

## 2.2 Batch instructions

Follow the process steps for each batch of samples:

- Prepare fresh digest of Reference Lysate with undigested samples.

- Add Reference Lysate Digest to a Capture Plate from the kit.

- Perform Immunoaffinity Enrichment and LC-MS analysis.

- Repeat above steps on another day.

# Cell culture

| Cell line | ATCC Catalog Number | Culture Properties | Culture Conditions | Culture Method Used |
| --- | --- | --- | --- | --- |
| T-47D | HTB-133 | adherent | 1 | A |
| CCRF-CEM | CCL-119 | suspension | 1 | A |
| COLO205 | CCL-222 | mixed ad-herent - su-spension | 1 | A |
| COR-L23 |  | adherent | 1 | A |
| NCI-H2444 | CRL-5945 | adherent | 1 | A |
| NCI-H2122 | CRL-5985 | adherent | 1 | A |
| NCI-H1792 | CRL-5895 | adherent | 1 | A |
| HEPG2 | HB-8065 | adherent | 1 | B |
| NCI H226 | CRL-5826 | adherent | 1 | A |
| K-562 | CCL-243 | suspension | 1 | A |
| mCF10a-EV | James A. Wells (UCSF) | adherent | 1 | D |
| mCF10a-KRAS | James A. Wells (UCSF) | adherent | 1 | D |
| A375 | CRL-1619 | adherent | 1 | P |
| SK-MEL-2 | HTB-68 | adherent | 1 | B |
| HT-29 | HTB-38 | adherent | 1 | G |
| HCT 116 | CCL-247 | adherent | 1 | G |
| A549 | CCL-185 | adherent | 1 | A |
| BxPc-3 | CRL-1687 | adherent | 1 | A |

| Culture Conditions |  |
| --- | --- |
| 1. 37°C, 95% air; 5% carbon dioxide |  |
| Culture Methods |  |
| A. 500 mL RPMI 1640 (Life Tech cat no. 11875093) |  |
| 50 mL HI FBS (Life Tech cat no. 16140071) |  |
| 5 mL L-Glutamine (Lonza cat no BW17-605E) |  |
| B. 500 mL EMEM (ATCC cat no. 30-2003) |  |
| 50 mL HI FBS (Life Tech cat no. 16140071) |  |
| 5 mL L-Glutamine (Lonza cat no BW17-605E) |  |
| D. Growth Medium. Component; Stock Conc.; Volume to add; Final Conc.. DMEM/F-121; 500 mL. Horse Serum2; 25 mL; 5%. Pen-Strep3; 100X; 5 mL; 1%. EGF4; 20 µg/mL; 0.5 mL; 20 ng/mL. Insulin5; 10 mg/mL; 0.5 mL; 10 µg/mL. Hydrocortisone6; 0.5 mg/mL; 0.5 mL; 0.5 µg/mL. Cholera Toxin7; 100 µg/mL; 0.5 mL; 100 ng/mL. |  |

Notes:

1) DMEM/F-12 (Gibco #11320): Dulbecco’s Modified Eagle’s Medium (DMEM)/Ham’s Nutrient Mixture F12 (1:1).

2) Horse Serum (Gibco #16050): 25 mL aliquots stored at -20°C. Thaw aliquot in 37°C water bath, then sterile filter before adding to DMEM/F12 media. Add 25 mL filtered HS per 500 mL growth medium bottle = 5% final.

3) Penicillin-Streptomycin (Gibco #15140-122): Penicillin at 10,000 U/mL and Streptomycin at 10,000 ug/mL.

4) EGF (PeproTech #AF-100-15): Resuspend at 20 ug/mL in sterile H2O. Sterile filter through a 0.2 um filter and store aliquots at -20°C. Add 0.5 mL EGF at 20 ug/mL per 500 mL growth medium bottle, resulting in 20 ng/mL final.

5) Insulin (Sigma #I6634): Resuspend at 10 mg/mL in sterile H2O containing 1% glacial acetic acid. Shake solution and allow 10-15 min to reconstitute. Sterile filter through a 0.2 ìm filter and store aliquots at -20°C. Add 0.5 mL Insulin at 10 mg/mL per 500 mL growth medium bottle, resulting in 10 ug/mL final.

6) Hydrocortisone (Sigma #H088): Resuspend at 0.5 mg/mL in 200 proof ethanol. Sterile filter through a 0.2 um filter and store aliquots at -20°C. Add 0.5 mL Hydrocortisone at 0.5 mg/mL per 500 mL growth medium bottle, resulting in 5 ug/mL final.

7) Cholera Toxin (Sigma #C8052): Resuspend at 100 ug/mL in sterile H2O and allow about 10 min to reconstitute. Sterile filter through a 0.2 um filter and store aliquots at 4°C. Add 0.5 mL Cholera Toxin at 100 ug/mL per 500 mL growth medium bottle, resulting in 100 ng/mL final.

G. 500 mL McCoy's 5A (ATCC cat no. 30-2007)

50 mL HI FBS (Life Tech cat no. 16140071)

5 mL L-Glutamine (Lonza cat no BW17-605E)

P. 500 mL DMEM (ATCC cat no. 30-2002)

50 mL HI FBS (Life Tech cat no. 16140071)

5 mL L-Glutamine (Lonza cat no BW17-605E)

# Antibody beads crosslinking

| Required Solutions | Description |
| --- | --- |
| Crosslinking solution | 20 mM DMP in 200 mM triethanolamine, pH 8.5 (made fresh) |
| Quench Solution | 150 mM monoethanolamine, pH 9.0 |
| Wash Solution | 1x PBS + 0.01% CHAPS |
| Resuspension Solution | 1x PBS + 0.01% CHAPS + 0.03% NaN3 |

Method:

- Wash magnetic protein G beads in Eppendorf tubes with 800 ul 1XPBS + 0.01% CHAPS.

- Put the tubes on a magnet and take out the supernatant.

- Wash magnetic protein G beads in Eppendorf tubes with 800 ul 1XPBS + 0.01% CHAPS.

- Put the tubes on a magnet and take out the supernatant.

- Add required amount of magnetic protein G beads to Eppendorf tubes, at a 5 ug mAb / 1 uL of beads ratio. Put the tubes on a magnet and take out the supernatant.

- Add required amout of antibody to each tube individually (either 1.6 mL, 2.0 mL, or 5.0 mL Eppendorf tubes will do depending on the concentration of the antibody and amount to be crosslinked).

- If needed, bring up total volume in each tube to a minimum amount for proper mixing using 1x PBS (0.25ml for 1.6 mL or 2.0 mL tubes, 2.0 mL for 5 mL tubes).

- Make sure all caps are closed tightly; vortex each tube to ensure all beads are in solution.

- Incubate at 4 oC overnight with end-over-end mixing.

- Put the tubes on a magnet and take out the supernatant.

- Add 400 ul fresh made 20mM DMP in 200 mM triethanolamine, pH 8.5 to beads.

- Incubate beads at room temperature for 30 minutes with end-over-end mixing.

- Put the tubes on a magnet and take out the supernatant.

- Add 400ul 150 mM monoethanolamine, pH 9.0 to beads.

- Incubate at Room temperature for 30 minutes with end-over-end mixing.

- Put the tubes on a magnet and take out the supernatant.

- Wash 800 ul with 1x PBS + 0.01% CHAPS

- Put the tubes on a magnet and take out the supernatant.

- Wash 800 ul with 1x PBS + 0.01% CHAPS

- Put the tubes on a magnet and take out the supernatant.

- Resuspense beads with required amount 1x PBS + 0.01% CHAPS + 0.03% NaN3 to final concentration of antibody at 2 ug/ul.

# Peptide Mastermix Preparation

Purpose

To have the same mixes at each site for determination of MRM parameters and analysis of samples.

Materials and Reagents

- Water, HPLC grade (Burdick & Jackson, AH365-4)

- Acetonitrile, HPLC grade (Burdick & Jackson, AH015-4)

- Methanol, HPLC grade (Burdick & Jackson, AH230-4)

- Formic Acid, Optima LC-MS Grade, ≥99.5%  (Fisher, A117-50)

- 1.5 mL Safe-Lock Tubes (Eppendorf, 022363204)

- 0.5 mL Safe-Lock Tubes (Eppendorf, 022363611)

- epT.I.P.S. reloads (Eppendorf)

- Repeater® M4 Pipette 1 µL – 10 mL (Eppendorf, 4982000322)

- Autosampler vials, 250 µL (Thermo, 160133)

- Pierce Retention Time Calibration Mixture (Thermo, 88321)

Reagent Preparation

- 30% ACN/ 0.1% Formic Acid

- 3% ACN/ 0.1% Formic Acid (A Solvent)

Peptide Master Mix Preparation Protocol

- Remove peptides from storage at -80° C and thawed on ice.

- Vortex and centrifuge each peptide stock solution for 20 – 30 s at 14k RPMs at 10° C.

- Sort peptide stocks into 4 groups: unmod light, unmod heavy, phos light, phos heavy.

- Combine 60 µL of each peptide into one tube per mix.

  - Unmodified, Light (6960 µL)

  - Unmodified, Heavy (7020 µL)

  - Phospho, Light (3000 µL)

  - Phospho, Heavy (2940 µL)

- Add an additional 60 µL aliquot of buffer (30% ACN, 0.1% FA) to Unmod, Light and Phospho, Heavy. Final concentrations:

  - Unmodified 1 & 2 – 8.55 µM

  - Phospho 3 & 4 – 20 µM

- Mix 10 µL of each light and heavy mix, resulting in concentrations of:

  - Unmod – 4.275 µM

  - Phospho – 10 µM

- Dilute unmodified Light + SIS Mix to 0.625 µM and phospho Light + SIS Mix to 2.5 µM.

  - Unmod: 14.62 µL of Light/SIS Mix (4.275 µM) + 85.38 µL A Solvent

  - Phospho: 10 µL of Light/SIS Mix (10 µM) + 30 µL A Solvent

- Make a mix with all four with unmod conc of 5 fmol/uL and phospho conc of 50 fmol/uL.

  - Pipette 486 µL of A solvent (3% ACN, 0.1% FA) into a fresh tube.

  - Add 4 µL of unmod dilution made in step 4.

  - Add 10 µL of phospho dilution made in step 4.

  - Add 5 µL of 500 fmol/ µL PRTC peptides for system suitability standards and retention time alignment (iRT).

  - Vortex to mix then pipet 100 µL into a labeled autosampler vial.

| MM1 |  |  |  | MM2 |  |  |  | MM3 |  |  |  | MM4 |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Peptide | Prot |  | Type | Peptide | Protein |  | Type | Peptide | Protein | Type | Only | Peptide | Protein | Type |
| FFAGIVWQHVYEK | AKT1 | L |  | FFAGIVWQHVYEK | AKT1 | SIS |  | (pT)FCGTPEYLAPEVLEDNDYGR | AKT1 | L |  | (pS)PDLAPTPAPQSTPR | EIF2A | SIS |
| FYGAEIVSALDYLHSEK | AKT1 | L |  | FYGAEIVSALDYLHSEK | AKT1 | SIS |  | RPHFPQF(pS)YSASGTA | AKT1 | L |  | (pT)FCGTPEYLAPEVLEDNDYGR | AKT1 | SIS |
| RPHFPQFSYSASGTA | AKT1 | L |  | RPHFPQFSYSASGTA | AKT1 | SIS |  | LN(pT)SDFQK | AKT1S1 | L |  | (pY)SSDPTGALTEDSIDDTFLPVPEYINQSVPK | EGFR | SIS |
| SLLSGLLK | AKT1 | L |  | SLLSGLLK | AKT1 | SIS |  | THFPQF(pS)YSASIR | AKT2 | L |  | AA(pS)MDNNSK | FOXO1 | SIS |
| TFCGTPEYLAPEVLEDNDYGR | AKT1 | L |  | TFCGTPEYLAPEVLEDNDYGR | AKT1 | SIS |  | THFPQF(pS)YSASIRE | AKT2 | L |  | AV(pS)MDNSNK | FOXO3 | SIS |
| LNTSDFQK | AKT1S1 | L |  | LNTSDFQK | AKT1S1 | SIS |  | RPHFPQF(pS)YSASGR | AKT3 | L |  | AV(pS)MDNSNKYTK | FOXO3 | SIS |
| EGISDGATMK | AKT2 | L |  | EGISDGATMK | AKT2 | SIS |  | RPHFPQF(pS)YSASGRE | AKT3 | L |  | D(pS)GYYWEVPPSEVQLLK | ARAF | SIS |
| FYGAEIVSALEYLHSR | AKT2 | L |  | FYGAEIVSALEYLHSR | AKT2 | SIS |  | D(pS)GYYWEVPPSEVQLLK | ARAF | L |  | D(pS)SDDWEIPDGQITVGQR | BRAF | SIS |
| LLPPFKPQVTSEVDTR | AKT2 | L |  | LLPPFKPQVTSEVDTR | AKT2 | SIS |  | NLGYRD(pS)GYYWEVPPSEVQLLK | ARAF | L |  | FI(camC)VTPTT(camC)SNTIDLPM(pS)PR | STAT3 | SIS |
| SLLAGLLK | AKT2 | L |  | SLLAGLLK | AKT2 | SIS |  | D(pS)SDDWEIPDGQITVGQR | BRAF | L |  | GDGGSTTGLSA(pT)PPASLPGSLTNVK | BRAF | SIS |
| THFPQFSYSASIR | AKT2 | L |  | THFPQFSYSASIR | AKT2 | SIS |  | GDGGSTTGLSA(pT)PPASLPGSLTNVK | BRAF | L |  | GF(pS)FVATGLMEDDGKPR | RPS6KA1 | SIS |
| THFPQFSYSASIRE | AKT2 | L |  | THFPQFSYSASIRE | AKT2 | SIS |  | RD(pS)SDDWEIPDGQITVGQR | BRAF | L |  | GSHQISLDNPD(pY)QQDFFPK | EGFR | SIS |
| DEVAHTLTESR | AKT3 | L |  | DEVAHTLTESR | AKT3 | SIS |  | RT(pS)MGGTQQQFVEGVR | CTNNB1 | L |  | GSTAENAE(pY)LR | EGFR | SIS |
| EGITDAATMK | AKT3 | L |  | EGITDAATMK | AKT3 | SIS |  | T(pS)MGGTQQQFVEGVR | CTNNB1 | L |  | GTPTAENPE(pY)LGLDVPV | ERBB2 | SIS |
| FYGAEIVSALDYLHSGK | AKT3 | L |  | FYGAEIVSALDYLHSGK | AKT3 | SIS |  | (pY)SSDPTGALTEDSIDDTFLPVPEYINQSVPK | EGFR | L |  | HTDDEM(pT)G(pY)VATR | MAPK14 | SIS |
| RPHFPQFSYSASGR | AKT3 | L |  | RPHFPQFSYSASGR | AKT3 | SIS |  | GSHQISLDNPD(pY)QQDFFPK | EGFR | L |  | HTDDEM(pT)GYVATR | MAPK14 | SIS |
| RPHFPQFSYSASGRE | AKT3 | L |  | RPHFPQFSYSASGRE | AKT3 | SIS |  | GSTAENAE(pY)LR | EGFR | L |  | HTDDEMTG(pY)VATR | MAPK14 | SIS |
| TDGSFIGYK | AKT3 | L |  | TDGSFIGYK | AKT3 | SIS |  | Y(pS)(pS)DPTGALTEDSIDDTFLPVPEYINQSVPK | EGFR | L | * | IADPEHDHTGFL(pT)E(pY)VATR | MAPK3 | SIS |
| DSGYYWEVPPSEVQLLK | ARAF | L |  | DSGYYWEVPPSEVQLLK | ARAF | SIS |  | YSSDPTGALTEDSIDDTFLPVPE(pY)INQSVPK | EGFR | L |  | IADPEHDHTGFL(pT)EYVATR | MAPK3 | SIS |
| GLNQDCCVVYR | ARAF | L |  | GLNQDCCVVYR | ARAF | SIS |  | (pS)PDLAPTPAPQSTPR | EIF2A | L |  | IADPEHDHTGFLTE(pY)VATR | MAPK3 | SIS |
| IGTGSFGTVFR | ARAF | L |  | IGTGSFGTVFR | ARAF | SIS |  | SDK(pS)PDLAPTPAPQSTPR | EIF2A | L |  | L(camC)DFGVSGQLID(pS)MANSFVGTR | MAP2K2 | SIS |
| NLGYRDSGYYWEVPPSEVQLLK | ARAF | L |  | NLGYRDSGYYWEVPPSEVQLLK | ARAF | SIS |  | SDKSPDLAPTPAPQS(pT)PR | EIF2A | L |  | LLDIDETE(pY)HADGGK | ERBB2 | SIS |
| TQADELPACLLSAAR | ARAF | L |  | TQADELPACLLSAAR | ARAF | SIS |  | SPDLAPTPAPQS(pT)PR | EIF2A | L |  | LN(pT)SDFQK | AKT1S1 | SIS |
| VSQPTAEQAQAFK | ARAF | L |  | VSQPTAEQAQAFK | ARAF | SIS |  | GTPTAENPE(pY)LGLDVPV | ERBB2 | L |  | NLGYRD(pS)GYYWEVPPSEVQLLK | ARAF | SIS |
| AGFQTEDFSLYACASPK | BRAF | L |  | AGFQTEDFSLYACASPK | BRAF | SIS |  | LLDIDETE(pY)HADGGK | ERBB2 | L |  | RD(pS)SDDWEIPDGQITVGQR | BRAF | SIS |
| DSSDDWEIPDGQITVGQR | BRAF | L |  | DSSDDWEIPDGQITVGQR | BRAF | SIS |  | AA(pS)MDNNSK | FOXO1 | L |  | RPHFPQF(pS)YSASGR | AKT3 | SIS |
| GDGGSTTGLSATPPASLPGSLTNVK | BRAF | L |  | GDGGSTTGLSATPPASLPGSLTNVK | BRAF | SIS |  | AV(pS)MDNSNK | FOXO3 | L |  | RPHFPQF(pS)YSASGRE | AKT3 | SIS |
| GLIPECCAVYR | BRAF | L |  | GLIPECCAVYR | BRAF | SIS |  | AV(pS)MDNSNKYTK | FOXO3 | L |  | RPHFPQF(pS)YSASGTA | AKT1 | SIS |
| RDSSDDWEIPDGQITVGQR | BRAF | L |  | RDSSDDWEIPDGQITVGQR | BRAF | SIS |  | TT(pS)FAESCKPVQQPSAFGSMK | GSK3B | L |  | RT(pS)MGGTQQQFVEGVR | CTNNB1 | SIS |
| SNNIFLHEDLTVK | BRAF | L |  | SNNIFLHEDLTVK | BRAF | SIS |  | L(camC)DFGVSGQLID(pS)MANSFVGTR | MAP2K2 | L |  | SDK(pS)PDLAPTPAPQSTPR | EIF2A | SIS |
| ACQEQIEALLESSLR | CCND1 | L |  | ACQEQIEALLESSLR | CCND1 | SIS |  | VADPDHDHTGFL(pT)E(pY)VATR | MAPK1 | L |  | SDKSPDLAPTPAPQS(pT)PR | EIF2A | SIS |
| AEETCAPSVSYFK | CCND1 | L |  | AEETCAPSVSYFK | CCND1 | SIS |  | VADPDHDHTGFL(pT)EYVATR | MAPK1 | L |  | SPDLAPTPAPQS(pT)PR | EIF2A | SIS |
| AYPDANLLNDR | CCND1 | L |  | AYPDANLLNDR | CCND1 | SIS |  | VADPDHDHTGFLTE(pY)VATR | MAPK1 | L |  | ST(pS)TPNVHMVSTTLPVDSR | RAF1 | SIS |
| FLSLEPVKK | CCND1 | L |  | FLSLEPVKK | CCND1 | SIS |  | HTDDEM(pT)G(pY)VATR | MAPK14 | L |  | T(pS)MGGTQQQFVEGVR | CTNNB1 | SIS |
| VIKCDPDCLR | CCND1 | L |  | VIKCDPDCLR | CCND1 | SIS |  | HTDDEM(pT)GYVATR | MAPK14 | L |  | TAGTSFMM(pT)P(pY)VVTR | MAPK8 | SIS |
| GLDARPEVTR | CDH1 | L |  | GLDARPEVTR | CDH1 | SIS |  | HTDDEMTG(pY)VATR | MAPK14 | L |  | TAGTSFMM(pT)PYVVTR | MAPK8 | SIS |
| GQVPENEANVVITTLK | CDH1 | L |  | GQVPENEANVVITTLK | CDH1 | SIS |  | IADPEHDHTGFL(pT)E(pY)VATR | MAPK3 | L |  | TAGTSFMMTP(pY)VVTR | MAPK8 | SIS |
| NDVAPTLMSVPR | CDH1 | L |  | NDVAPTLMSVPR | CDH1 | SIS |  | IADPEHDHTGFL(pT)EYVATR | MAPK3 | L |  | TD(pS)YSAGQSVEILDGVELGEPAHK | MTOR | SIS |
| NTGVISVVTTGLDR | CDH1 | L |  | NTGVISVVTTGLDR | CDH1 | SIS |  | IADPEHDHTGFLTE(pY)VATR | MAPK3 | L |  | TGTTVPESIH(pS)FIGDGLVKPEALNK | MTOR | SIS |
| TAYFSLDTR | CDH1 | L |  | TAYFSLDTR | CDH1 | SIS |  | TAGTSFMM(pT)P(pY)VVTR | MAPK8 | L |  | TGTTVPESIH(pS)FIGDGLVKPEALNKK | MTOR | SIS |
| GPFPQELVR | CDH2 | L |  | GPFPQELVR | CDH2 | SIS |  | TAGTSFMM(pT)PYVVTR | MAPK8 | L |  | THFPQF(pS)YSASIR | AKT2 | SIS |
| IDPVNGQITTIAVLDR | CDH2 | L |  | IDPVNGQITTIAVLDR | CDH2 | SIS |  | TAGTSFMMTP(pY)VVTR | MAPK8 | L |  | THFPQF(pS)YSASIRE | AKT2 | SIS |
| LNGDFAQLNLK | CDH2 | L |  | LNGDFAQLNLK | CDH2 | SIS |  | TD(pS)YSAGQSVEILDGVELGEPAHK | MTOR | L |  | TRTD(pS)YSAGQSVEILDGVELGEPAHK | MTOR | SIS |
| LSDPANWLK | CDH2 | L |  | LSDPANWLK | CDH2 | SIS |  | TGTTVPESIH(pS)FIGDGLVKPEALNK | MTOR | L |  | TT(pS)FAESCKPVQQPSAFGSMK | GSK3B | SIS |
| SAAPHPGDIGDFINEGLK | CDH2 | L |  | SAAPHPGDIGDFINEGLK | CDH2 | SIS |  | TGTTVPESIH(pS)FIGDGLVKPEALNKK | MTOR | L |  | VADPDHDHTGFL(pT)E(pY)VATR | MAPK1 | SIS |
| RTSMGGTQQQFVEGVR | CTNNB1 | L |  | RTSMGGTQQQFVEGVR | CTNNB1 | SIS |  | TRTD(pS)YSAGQSVEILDGVELGEPAHK | MTOR | L |  | VADPDHDHTGFL(pT)EYVATR | MAPK1 | SIS |
| TSMGGTQQQFVEGVR | CTNNB1 | L |  | TSMGGTQQQFVEGVR | CTNNB1 | SIS |  | ST(pS)TPNVHMVSTTLPVDSR | RAF1 | L |  | VADPDHDHTGFLTE(pY)VATR | MAPK1 | SIS |
| GSHQISLDNPDYQQDFFPK | EGFR | L |  | GSHQISLDNPDYQQDFFPK | EGFR | SIS |  | GF(pS)FVATGLMEDDGKPR | RPS6KA1 | L |  | Y(camC)RPESQEHPEADPGSAAP(pY)LK | STAT3 | SIS |
| GSTAENAEYLR | EGFR | L |  | GSTAENAEYLR | EGFR | SIS |  | FI(camC)VTPTT(camC)SNTIDLPM(pS)PR | STAT3 | L |  | YSSDPTGALTEDSIDDTFLPVPE(pY)INQSVPK | EGFR | SIS |
| IPLENLQIIR | EGFR | L |  | IPLENLQIIR | EGFR | SIS |  | Y(camC)RPESQEHPEADPGSAAP(pY)LK | STAT3 | L |  |  |  |  |
| NLQEILHGAVR | EGFR | L |  | NLQEILHGAVR | EGFR | SIS |  |  |  |  |  |  |  |  |
| YSSDPTGALTEDSIDDTFLPVPEYINQSVPK | EGFR | L |  | YSSDPTGALTEDSIDDTFLPVPEYINQSVPK | EGFR | SIS |  |  |  |  |  |  |  |  |
| SDKSPDLAPTPAPQSTPR | EIF2A | L |  | SDKSPDLAPTPAPQSTPR | EIF2A | SIS |  |  |  |  |  |  |  |  |
| SPDLAPTPAPQSTPR | EIF2A | L |  | SPDLAPTPAPQSTPR | EIF2A | SIS |  |  |  |  |  |  |  |  |
| ELVSEFSR | ERBB2 | L |  | ELVSEFSR | ERBB2 | SIS |  |  |  |  |  |  |  |  |
| GIWIPDGENVK | ERBB2 | L |  | GIWIPDGENVK | ERBB2 | SIS |  |  |  |  |  |  |  |  |
| GTPTAENPEYLGLDVPV | ERBB2 | L |  | GTPTAENPEYLGLDVPV | ERBB2 | SIS |  |  |  |  |  |  |  |  |
| LLDIDETEYHADGGK | ERBB2 | L |  | LLDIDETEYHADGGK | ERBB2 | SIS |  |  |  |  |  |  |  |  |
| ANDALQVLGLLFSLAR | ERBB3 | L |  | ANDALQVLGLLFSLAR | ERBB3 | SIS |  |  |  |  |  |  |  |  |
| GDSAYHSQR | ERBB3 | L |  | GDSAYHSQR | ERBB3 | SIS |  |  |  |  |  |  |  |  |
| LTFQLEPNPHTK | ERBB3 | L |  | LTFQLEPNPHTK | ERBB3 | SIS |  |  |  |  |  |  |  |  |
| SLEATDSAFDNPDYWHSR | ERBB3 | L |  | SLEATDSAFDNPDYWHSR | ERBB3 | SIS |  |  |  |  |  |  |  |  |
| YLERGESIEPLDPSEK | ERBB3 | L |  | YLERGESIEPLDPSEK | ERBB3 | SIS |  |  |  |  |  |  |  |  |
| APHPFGVPAPSAGAYSR | FOS | L |  | APHPFGVPAPSAGAYSR | FOS | SIS |  |  |  |  |  |  |  |  |
| GSSSNEPSSDSLSSPTLLAL | FOS | L |  | GSSSNEPSSDSLSSPTLLAL | FOS | SIS |  |  |  |  |  |  |  |  |
| LEFILAAHR | FOS | L |  | LEFILAAHR | FOS | SIS |  |  |  |  |  |  |  |  |
| SALQTEIANLLK | FOS | L |  | SALQTEIANLLK | FOS | SIS |  |  |  |  |  |  |  |  |
| TEPFDDFLFPASSR | FOS | L |  | TEPFDDFLFPASSR | FOS | SIS |  |  |  |  |  |  |  |  |
| AVSMDNSNK | FOXO3 | L |  | AVSMDNSNK | FOXO3 | SIS |  |  |  |  |  |  |  |  |
| AVSMDNSNKYTK | FOXO3 | L |  | AVSMDNSNKYTK | FOXO3 | SIS |  |  |  |  |  |  |  |  |
| LLEYTPTAR | GSK3B | L |  | LLEYTPTAR | GSK3B | SIS |  |  |  |  |  |  |  |  |
| QTLPVIYVK | GSK3B | L |  | QTLPVIYVK | GSK3B | SIS |  |  |  |  |  |  |  |  |
| TPPEAIALCSR | GSK3B | L |  | TPPEAIALCSR | GSK3B | SIS |  |  |  |  |  |  |  |  |
| TTSFAESCKPVQQPSAFGSMK | GSK3B | L |  | TTSFAESCKPVQQPSAFGSMK | GSK3B | SIS |  |  |  |  |  |  |  |  |
| VIGNGSFGVVYQAK | GSK3B | L |  | VIGNGSFGVVYQAK | GSK3B | SIS |  |  |  |  |  |  |  |  |
| IPEQILGK | MAP2K1 | L |  | IPEQILGK | MAP2K1 | SIS |  |  |  |  |  |  |  |  |
| ISELGAGNGGVVFK | MAP2K1 | L |  | ISELGAGNGGVVFK | MAP2K1 | SIS |  |  |  |  |  |  |  |  |
| VSHKPSGLVMAR | MAP2K1 | L |  | VSHKPSGLVMAR | MAP2K1 | SIS |  |  |  |  |  |  |  |  |
| L(camC)DFGVSGQLIDSMANSFVGTR | MAP2K2 | L |  | L(camC)DFGVSGQLIDSMANSFVGTR | MAP2K2 | SIS |  |  |  |  |  |  |  |  |
| VADPDHDHTGFLTEYVATR | MAPK1 | L |  | VADPDHDHTGFLTEYVATR | MAPK1 | SIS |  |  |  |  |  |  |  |  |
| HTDDEMTGYVATR | MAPK14 | L |  | HTDDEMTGYVATR | MAPK14 | SIS |  |  |  |  |  |  |  |  |
| ALDLLDR | MAPK3 | L |  | ALDLLDR | MAPK3 | SIS |  |  |  |  |  |  |  |  |
| GQPFDVGPR | MAPK3 | L |  | GQPFDVGPR | MAPK3 | SIS |  |  |  |  |  |  |  |  |
| IADPEHDHTGFLTEYVATR | MAPK3 | L |  | IADPEHDHTGFLTEYVATR | MAPK3 | SIS |  |  |  |  |  |  |  |  |
| NYLQSLPSK | MAPK3 | L |  | NYLQSLPSK | MAPK3 | SIS |  |  |  |  |  |  |  |  |
| SQQLSNDHICYFLYQILR | MAPK3 | L |  | SQQLSNDHICYFLYQILR | MAPK3 | SIS |  |  |  |  |  |  |  |  |
| TAGTSFMMTPYVVTR | MAPK8 | L |  | TAGTSFMMTPYVVTR | MAPK8 | SIS |  |  |  |  |  |  |  |  |
| DLELAVPGTYDPNQPIIR | MTOR | L |  | DLELAVPGTYDPNQPIIR | MTOR | SIS |  |  |  |  |  |  |  |  |
| IQSIAPSLQVITSK | MTOR | L |  | IQSIAPSLQVITSK | MTOR | SIS |  |  |  |  |  |  |  |  |
| LFDAPEAPLPSR | MTOR | L |  | LFDAPEAPLPSR | MTOR | SIS |  |  |  |  |  |  |  |  |
| LTESLDFTDYASR | MTOR | L |  | LTESLDFTDYASR | MTOR | SIS |  |  |  |  |  |  |  |  |
| TDSYSAGQSVEILDGVELGEPAHK | MTOR | L |  | TDSYSAGQSVEILDGVELGEPAHK | MTOR | SIS |  |  |  |  |  |  |  |  |
| TGTTVPESIHSFIGDGLVKPEALNK | MTOR | L |  | TGTTVPESIHSFIGDGLVKPEALNK | MTOR | SIS |  |  |  |  |  |  |  |  |
| TGTTVPESIHSFIGDGLVKPEALNKK | MTOR | L |  | TGTTVPESIHSFIGDGLVKPEALNKK | MTOR | SIS |  |  |  |  |  |  |  |  |
| TRTDSYSAGQSVEILDGVELGEPAHK | MTOR | L |  | TRTDSYSAGQSVEILDGVELGEPAHK | MTOR | SIS |  |  |  |  |  |  |  |  |
| VLGLLGALDPYK | MTOR | L |  | VLGLLGALDPYK | MTOR | SIS |  |  |  |  |  |  |  |  |
| AQEALDFYGEVR | PTEN | L |  | AQEALDFYGEVR | PTEN | SIS |  |  |  |  |  |  |  |  |
| GVTIPSQR | PTEN | L |  | GVTIPSQR | PTEN | SIS |  |  |  |  |  |  |  |  |
| IYNLCAER | PTEN | L |  | IYNLCAER | PTEN | SIS |  |  |  |  |  |  |  |  |
| IYSSNSGPTR | PTEN | L |  | IYSSNSGPTR | PTEN | SIS |  |  |  |  |  |  |  |  |
| NHLDYRPVALLFHK | PTEN | L |  | NHLDYRPVALLFHK | PTEN | SIS |  |  |  |  |  |  |  |  |
| DSSYYWEIEASEVMLSTR | RAF1 | L |  | DSSYYWEIEASEVMLSTR | RAF1 | SIS |  |  |  |  |  |  |  |  |
| GLQPECCAVFR | RAF1 | L |  | GLQPECCAVFR | RAF1 | SIS |  |  |  |  |  |  |  |  |
| GYASPDLSK | RAF1 | L |  | GYASPDLSK | RAF1 | SIS |  |  |  |  |  |  |  |  |
| STSTPNVHMVSTTLPVDSR | RAF1 | L |  | STSTPNVHMVSTTLPVDSR | RAF1 | SIS |  |  |  |  |  |  |  |  |
| VVDPTPEQFQAFR | RAF1 | L |  | VVDPTPEQFQAFR | RAF1 | SIS |  |  |  |  |  |  |  |  |
| LLYGEPPKSPR | RASGRF1 | L |  | LLYGEPPKSPR | RASGRF1 | SIS |  |  |  |  |  |  |  |  |
| NSLDYAK | RASGRF1 | L |  | NSLDYAK | RASGRF1 | SIS |  |  |  |  |  |  |  |  |
| SLELLFASGQNNK | RASGRF1 | L |  | SLELLFASGQNNK | RASGRF1 | SIS |  |  |  |  |  |  |  |  |
| GFSFVATGLMEDDGKPR | RPS6KA1 | L |  | GFSFVATGLMEDDGKPR | RPS6KA1 | SIS |  |  |  |  |  |  |  |  |
| ALETIGANLQK | RPTOR | L |  | ALETIGANLQK | RPTOR | SIS |  |  |  |  |  |  |  |  |
| SLIVAGLGDGSIR | RPTOR | L |  | SLIVAGLGDGSIR | RPTOR | SIS |  |  |  |  |  |  |  |  |
| SYNCTPVSSPR | RPTOR | L |  | SYNCTPVSSPR | RPTOR | SIS |  |  |  |  |  |  |  |  |
| VLDTSSLTQSAPASPTNK | RPTOR | L |  | VLDTSSLTQSAPASPTNK | RPTOR | SIS |  |  |  |  |  |  |  |  |
| VLNSIAYK | RPTOR | L |  | VLNSIAYK | RPTOR | SIS |  |  |  |  |  |  |  |  |
| FI(camC)VTPTT(camC)SNTIDLPMSPR | STAT3 | L |  | FI(camC)VTPTT(camC)SNTIDLPMSPR | STAT3 | SIS |  |  |  |  |  |  |  |  |
| Y(camC)RPESQEHPEADPGSAAPYLK | STAT3 | L |  | Y(camC)RPESQEHPEADPGSAAPYLK | STAT3 | SIS |  |  |  |  |  |  |  |  |

| MM1* and MM2* |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Protein | Peptide | mAb Mix | Peptide Mix | Recommended Fold Increase | Concentration (µM) |
| AKT1 | FFAGIVWQHVYEK | unmod | unmod | 10x | 20 |
| AKT2 | LLPPFKPQVTSEVDTR | unmod | unmod | 2x | 4 |
| ARAF | DSGYYWEVPPSEVQLLK | phos | unmod | 2x | 4 |
| ARAF | NLGYRDSGYYWEVPPSEVQLLK | phos | unmod | 2x | 4 |
| CCND1 | ACQEQIEALLESSLR | unmod | unmod | 2x | 4 |
| CCND1 | FLSLEPVKK | unmod | unmod | 10x | 20 |
| CDH2 | SAAPHPGDIGDFINEGLK | unmod | unmod | 10x | 20 |
| EGFR | GSHQISLDNPDYQQDFFPK | phos | unmod | 2x | 4 |
| EGFR | YSSDPTGALTEDSIDDTFLPVPEYINQSVPK | phos | unmod | 10x | 20 |
| ERBB2 | GTPTAENPEYLGLDVPV | phos | unmod | 2x | 4 |
| FOS | GSSSNEPSSDSLSSPTLLAL | unmod | unmod | 2x | 4 |
| MK01 | SQQLSNDHI(camC)YFLYQILR | unmod | unmod | 10x | 20 |
| MTOR | TGTTVPESIHSFIGDGLVKPEALNKK | phos | unmod | 2x | 4 |

| MM3* and MM4* |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Protein | Peptide | mAb Mix | Peptide Mix | Recommended Fold Increase | Concentration (µM) |
| AKT1 | (pT)FCGTPEYLAPEVLEDNDYGR | phos | phos | 10x | 20 |
| AKT2 | THFPQF(pS)YSASIRE | phos | phos | 2x | 4 |
| ARAF | D(pS)GYYWEVPPSEVQLLK | phos | phos | 10x | 20 |
| ARAF | NLGYRD(pS)GYYWEVPPSEVQLLK | phos | phos | 2x | 4 |
| EGFR | GSHQISLDNPD(pY)QQDFFPK | phos | phos | 2x | 4 |
| EGFR | (pY)SSDPTGALTEDSIDDTFLPVPEYINQSVPK | phos | phos | 25x | 50 |
| EGFR | YSSDPTGALTEDSIDDTFLPVPE(pY)INQSVPK | phos | phos | 25x | 50 |
| EGFR | Y(pS)(pS)DPTGALTEDSIDDTFLPVPEYINQSVPK | phos | phos | 25x | 50 |
| ERBB2 | GTPTAENPE(pY)LGLDVPV | phos | phos | 10x | 20 |
| FOXO1 | AA(pS)MDNNSK | phos | phos | 2x | 4 |
| FOXO3 | AV(pS)MDNSNK | phos | phos | 2x | 4 |
| MP2K1 | L(camC)DFGVSGQLID(pS)MANSFVGTR | phos | phos | 2x | 4 |
| MTOR | TD(pS)YSAGQSVEILDGVELGEPAHK | phos | phos | 10x | 20 |
| MTOR | TGTTVPESIH(pS)FIGDGLVKPEALNKK | phos | phos | 2x | 4 |
| MTOR | TRTD(pS)YSAGQSVEILDGVELGEPAHK | phos | phos | 10x | 20 |
| RAF1 | ST(pS)TPNVHMVSTTLPVDSR | phos | phos | 2x | 4 |

Analysis of Peptide Mixes by LC-MS/MS

Materials and Reagents

| Mobile phase A | Mobile phase B | LC autosampler | LC | MS | SS mix |
| --- | --- | --- | --- | --- | --- |
| 2% ACN, 0.1% FA | 90% ACN, 0.1% FA | Dionex Ultimate 3000, RS Autosampler (Nano) | Dionex Ultimate 3000, RSLCnano System | Thermo QExactive Plus | Pierce Retention Time Calibration Mixture |

- Prepare a new transfer vial with fresh A solvent (3% ACN, 0.1% FA) and two autosampler wash vials: one with A solvent and one with 50% methanol/40% ACN (wash solvent).

- Inject 5 µL (25 fmol of unmodified and 250 fmol of phospho) on QE plus (Thermo) using an 80 minute method (60 minute gradient) to acquire LC-MS/MS data.

LC-MS/MS Parameters

| Trap column | Acclaim PepmapTM 100, 100 µm × 2 cm, C18, 5 µm, 100A (Thermo) |
| --- | --- |
| Analytical column | Acclaim PepmapTM RSLC, 75 µm × 25 cm, C18, 2 µm, 100A (Thermo) |
| MS Resolution | MS1 – 70k, MS2 – 17.5k |
| MS Spray Conditions | 2400 V, Transfer tube 250° C |
| Scan Range | 380 – 1400 m/z |
| Full MS AGC/Max IT | 3e6, 200 ms |
| MS/MS AGC/Max IT | 1e5, 75 ms, TOP 12 |
| Min AGC | 6e2 |
| Dynamic Exclusion | 15s, 10 ppm |
| Isolation Window/Offset | 1.2 m/z, 0.2 m/z |

Gradient

| Time (min) | % B | % A |
| --- | --- | --- |
| 0 | 2 | 98 |
| 6 | 2 | 98 |
| 8 | 10 | 90 |
| 50 | 28 | 72 |
| 55 | 50 | 50 |
| 56.5 | 92 | 8 |
| 62.5 | 95 | 5 |
| 63.5 | 2 | 98 |
| 80 | 2 | 98 |

- Analyze results in Skyline. Check PRTC peptide stability and Light/SIS Ratios, which should equal 1.

- Adjust peptides with ratios of ± 20%.

- Repeat preparation of diluted peptides in buffer, as described above (25 fmol unmodified, 250 fmol phosphorylated).

- Aliquot 50 µL per tube for unmodified mixes (MM1 & MM2) and 25 µL per tube of phospho peptide mixes (MM3 & MM4). Store at -80o C until ready to use.

Analysis of Samples by LC-MRM

Materials and Reagents

| Mobile phase A | Mobile phase B | LC autosampler | LC | MS | SS mix |
| --- | --- | --- | --- | --- | --- |
| 2% ACN, 0.1% FA | 90% ACN, 0.1% FA | Dionex Ultimate 3000, RS Autosampler (Nano) | Dionex Ultimate 3000, RSLCnano System | Thermo TSQ Quantiva | Pierce Retention Time Calibration Mixture |

Analysis Procedure

- Prepare MS system suitability standards (PRTC) and blank injections (3% ACN, 0.1% FA) in labeled autosampler vials.

- Inject 5 µL of 5 fmol/µL PRTC peptides using the gradient optimized for RAS iMRM targets (see below).

- Set up scheduling using PRTC LC-MRM run and Skyline iRT library containing all RAS iMRM peptides.

LC-MRM Parameters

| Trap column | Acclaim PepmapTM 100, 100 µm × 2 cm, C18, 5 µm, 100A (Thermo) |
| --- | --- |
| Analytical column | Acclaim PepmapTM RSLC, 75 µm × 25 cm, C18, 2 µm, 100A (Thermo) |
| MS Resolution | Q1: 0.4, Q3: 0.7 |
| MS Spray Conditions | 2200 V, Transfer tube 275° C |
| CID Gas | 2 mTorr |
| Dwell time | 15 ms |

Gradient

# Generation of cell lysates

Protein lysates from cultured cells for Mass Spec.

Purpose:  Generate mass spec compatible protein lysates from cultured cells.

Background:	This protocol assumes:

- Cells have been harvested and washed in DPBS.

A.  Cell Lysates

- Preparation:

  - Turn on refrigerated micro-centrifuge and cool to 4oC.

  - Turn on Coulter Counter and prime aperture.

  - Thaw Phosphatase inhibitors.

  - Label and pre-cool 50 mL tubes.

  - Label and pre-cool micro-centrifuge tubes.

  - Label and pre-cool Cryo-vials.

  - Make fresh urea lysis buffer- see solution section below.

- Place tube with cell pellet on ice and add lysis buffer to a final concentration of 5x107 cells / mL.

- Resuspend cells in lysis buffer by dragging tube along a microfuge tube rack- do not pipette.

- Sonicate cells 2 x 12 sec. (550 Sonic Dimembrator, Fisher Scientific; knob set to 5)

  - Wipe down probe with water and ethanol between samples.

  - Place lysate on ice for ~20 sec. between sonications.

- Transfer lysate by pipette tip to micro-centrifuge tube, vortex 15 sec..

- Micro-centrifuge: 20k x g (14K RPM or full speed) / 10 min. / 4oC.

- Transfer supernatant to 1.0 mL cryo-vial

  - Note: if storing aliquots of a lysate, first transfer the lysate to a fresh micro-centrifuge tube to ensure homogeneous mixing of the lysate before aliquoting.

- Store lysates in -80oC.

- Determine protein concentration by BCA.

Solutions and Reagents:

Solutions:

- Lysis Buffer. Must be made fresh daily:

  - 4 Parts 7.5 M Urea (see below)

  - 1 Part  5x Lysis Buffer Stock Solution (see below).

  - Add 1% Sigma phosphatase cocktail 1

  - Add 1% Sigma phosphatase cocktail 2

  - Add 1% Sigma Protease Inhibitor

  - Mix well.

- 5x Lysis Buffer Stock Solution. May be made in advance and stored at room temp.

To Make 100 mL:

  - 12.5	mL  1M Tris (pH8.0)

  - 1.0 	mL  0.5 M EDTA

  - 1.0  	mL  0.5 M EGTA

  - HPLC water to 100 mL

  - Sterilize with 0.22 um filter.

- 7.5 M Urea. Make fresh daily.

  - Add 6.5 mL HPLC water to a 15 mL Falcon tube.

  - Add 4.50 g Urea to the 15 mL Falcon tube.

  - Mix until Urea is in solution, final volume should be 10 mL.

Reagents:

- Urea: Sigma Ultra Cat# U0631

- 1 M Tris (pH8.0): Sigma Cat# T2194

- EDTA: Sigma Cat# E7889

- EGTA: Sigma Cat# E0396

- TCEP: Sigma Cat# C4706

- HPLC water: Fisher Cat# W6-4

- Sigma Protease Inhibitor Cat# P8340

- Sigma Phosphatase Cocktail 2 Cat# P5726

- Sigma Phosphatase Cocktail 3 Cat# P0044

Final Urea buffer: 6M Urea, 25 mM Tris (pH8.0), 1 mM EDTA, 1 mM EGTA

# Protein Digestion

## 7.1 Materials and Reagents

- Lys-C (Endoproteinase Lys-C, Wako No. 129-02541)

- Trypsin, sequencing grade, (Promega, V5113)

- Urea, (Sigma, U0631)

- 0.5 M TCEP, (Pierce, 77720)

- Iodoacetamide (IAM), (Sigma, A3221-10VL) (note these are single-use vials. Alternative to weigh out amounts from another source.)

- 1 M Tris (pH8.0), (Sigma, T2694)

- EDTA (Sigma, E7889)

- EGTA (Sigma, E0396)

- Water, HPLC grade (Fisher, W5-1)

- Acetonitrile, HPLC grade (Fisher, A998-1)

- Formic Acid (EDM, 11670-1)

- Sigma Phosphatase Inhibitor Cocktail 2 (Sigma, P5726)

- Sigma Phosphatase Inhibitor Cocktail 3 (Sigma, P0044)

- 2 mL polypropylene 96 well deep well plates (ThermoFisher Scientific, 95040450) ; used for digestion

- XPierce pre-cut pierce-able films (Genesee Scientific, 12-140), or equivalent ; for digestions

- Thermomixer, or equivalent

## 7.2 Reagent Preparation

- 5 × Lysis Buffer Stock Solution. May be made in advance and stored at room temp.

  - 12.5 mL 1M Tris (pH 8.0).

  - 1.0 mL 0.5 M EDTA.

  - 1.0 mL 0.5 M EGTA.

  - Add HPLC water to 100 mL.

  - Filter with 0.22 μm filter.

- 7.5 M Urea. Must be made fresh daily:

  - Add 4.5 g Urea to a 15 mL Falcon tube.

  - Add 6 mL HPLC water and mix until Urea is in solution.

  - Add HPLC water to a final volume of 10 mL.

- Lysis Buffer. Must be made fresh daily:

  - 4 Parts 7.5 M Urea (see above).

  - 1 Part 5 × Lysis Buffer Stock Solution (see above).

  - Add 1% Sigma phosphatase cocktail 2.

  - Add 1% Sigma phosphatase cocktail 3.

  - Mix well.

- 0.2 M Tris, pH 8.0:

  - 4 parts HPLC water.

  - 1 part 1M Tris, pH 8.0.

- 0.5 M Iodoacetamide (IAM) Stock Solution.  Prepare immediately before use and keep out of light:

  - To one 56 mg vial of iodoacetamide, add 605 μL of Tris buffer (D above).

  - Mix until dissolved. Visually inspect vial to confirm fully solvated.

- Lys-C (Assuming 8 mg in 10 AU):

  - Prepare 50 mM Acetic Acid by mixing:

    - 17.8 mL HPLC water.

    - 7.2 mL 1% solution of Acetic Acid (HOAc, Sigma cat# 242853, stock conc. 17.4 M)

  - Dissolve 10 AU Lys-C in 6 mL 50 mM HOAc.

  - Transfer dissolved enzyme to a 50 mL tube.

  - Add 10.0 mL 50 mM HOAc for a final volume of 16 mL and an estimated conc. of 500 μg / mL.

  - Aliquot 0.5 mL / 1.7 mL microfuge tube (~250 ug / tube)

  - Store aliquots at -80 °C.

- Trypsin:

  - 100 μg resuspended in 200 μL of 0.2 M Tris pH 8.

- Quench:

  - 4 parts HPLC water.

  - 1 part formic acid.

- SIS peptide mastermix working solution

  - yeast reference digest: bulk yeast (S. cerevisiae) lysate digested with trypsin, desalted, and stored frozen in 3% acetonitrile / 0.1% formic acid at 30 μg /mL.

  - Working MM2: Add the following volumes of peptide mastermix to a final volume of 20 μL in yeast digest.

    - 2.5 μL MM2

    - 10 μL MM2*

  - Working MM4: Add the following volumes of peptide mastermix to a final volume of 40 μL in yeast digest.

    - 2.12 μL MM4

    - 20 μL MM4*

  - Working mastermix (prepare fresh daily): Add the following volumes of working MM2/4 (from steps 2 and 3 above) to a final volume of 1000 μL in yeast digest.

    - 2.0 μL working MM2

    - 5.0 μL working MM4

## 7.3 Digestion Protocol

|  | - Dilute the cell lysate to 2 mg/mL with lysis buffer. (note: supplied Reference lysate concentration is 7.15 mg/mL, see page 1 or page 11) |
| --- | --- |
|  | - Transfer 250 μL of cell lysate to a 2 mL deepwell plate. |
|  | - Place X-Pierce film on sample plate. |
|  | - Add 16.5 μL of 0.5 M TCEP. |
|  | - Incubate with mixing at 600 rpm on a Thermomixer for 30 min at 37 °C. |
|  | - Add 36 μL of 0.5 M IAM in Lysis buffer. |
|  | - Alkylate at room temperature for 30 min in the dark. |
|  | - Add 1200 μL of 0.2 M Tris, pH 8.0 to each tube to decrease the urea concentration to ~0.6 M. |
|  | - Add 20 μL LysC to each digest to achieve a 1:50 enzyme-to-substrate ratio. |
|  | - Incubate with mixing at 600 rpm for 2 h at 37 °C. |
|  | - Add 20 μL trypsin to each digest to achieve a 1:50 enzyme-to-substrate ratio. |
|  | - Incubate with mixing at 600 rpm for 16 h at 37 °C. |
|  | - Add 76.5 μL of 20% FA to each digest to quench the digestion for a final acid concentration of ~1%. |
|  | - Add 10 μL of heavy SIS mastermix working solution (I.4 above) to each sample. |

# Immunoaffinity Enrichment

[Note: this procedure is intended for sequential immunoaffinity enrichment. Each sample will be captured overnight twice, first with the anti-phosphorylated peptide mAb mixture (Ab beads multimix 1), then the remaining flowthrough (supernatant) will be subsequently captured with the anti-unmodified peptide mAb mixture (Ab beads multimix 2).]

## 8.1 Materials and Reagents

- 1 M Tris (pH 8.0), (Sigma, T2694)

- Water, HPLC grade (Fisher, W5-1)

- Acetonitrile, HPLC grade (Fisher, A998-1)

- Formic Acid (EDM, 11670-1)

- 10 × PBS (ThermoFisher Scientific, BP39920)

- CHAPS (Thermo Scientific, 28300)

- Acetic Acid, ACS reagent, ≥99.7% (Sigma, 242853)

- Citric Acid (Fluka, Sigma, 94068)

- KingFisher Magnetic Particle Processor

- KingFisher, PCR 96 well head

- Labquake tube rotator or shaker

- DynaMagTM-96 Side Skirted Magnet (12027, Life Technologies)

- 200 µL 96 well plates (ThermoFisher Scientific, 97002540) ; used for captures and washes

- Hard-Shell® Low-Profile Thin-Wall 96-Well Skirted PCR Plates (BioRad, HSP-9601) ; used for elution plate

- Square Matrix CapMats For 2 mL blocks (ThermoFisher Scientific, 4421) ; a sealer for capture plates

- AdvantageTM Purple Framed 96 Square Well Pattern Teflon Sealing Film (Analytical, 961201) ; alternative sealer for capture plates

- ThermalSeal® PCR Sealing Film

- Chemically resistant sealing foil (BioExpress, T-3025-8B), or equivalent ; for freezer storage

- 96-well Silicone Sealing Mat (Axygen Scientific, 22-513), or equivalent ; for autosampler

- Antibody-beads mastermix #1 solution (mAbs to phosphorylated peptide targets)

- Antibody-beads mastermix #2 solution (mAbs to unmodified peptide targets)

## 8.2 Reagent Preparation

- 1 × PBS / 0.01% CHAPS. May be made in advance and stored at room temp.

  - 5 mg CHAPS in 100mL LCMS-grade water.

  - Add 100 mL of 10 × PBS and 2 mL of 5% CHAPS to LCMS-grade water, final volume 1L.

- 0.1 × PBS / 0.01% CHAPS. May be made in advance and stored at room temp.

  - Add 10 mL of 10 × PBS and 2mL of 5% CHAPS to LCMS-grade water, final volume 1L.

- 5% acetic acid / 3% acetonitrile / 50 mM citrate:

  - Dissolve 8.4 g citric acid monohydrate in 200 mL of LCMS-grade water to make 200 mM citrate.

  - Mix 3.0 mL LCMS-grade acetonitrile, 5.0 mL acetic acid, 25 mL of 200 mM citrate, bring to 100 mL with water.

## 8.3 Peptide Immunoaffinity Enrichment Procedure

|  | - Resuspend the digested and desalted sample in 200 μL of 1 × PBS / 0.01% CHAPS using the 200 µL 96 well plates. If necessary, transfer the sample to the appropriate Capture Plate in the kit. |
| --- | --- |
|  | - Check that the pH is between 7.5 - 8.0. If necessary, adjust with 1 M Tris, pH 8.0 (5-10 µL should be sufficient). |
|  | - Mix the Antibody beads in the tube provided. Dispense enough beads to complete a capture plate into a separate tube. Add 24 µL of the Antibody beads multi-mix #1 (mAbs to phosphorylated peptide targets) to each sample well. When adding the beads, vortex the bead solution and pipet the master mix bead solution up and down several times to ensure well-mixed beads. |
|  | - Cover the plate with a Square Matrix CapMat (or Advantage Teflon Sealing Film). Align carefully and seal it by pressure to ensure that no liquid can leak from any of the wells. |
|  | - Incubate samples overnight at 4 °C with tumbling on a Labquake tube rotator or Lab Shaker. |
|  | - Remove samples from the tumbler and centrifuge at 800 × g for 30 seconds to remove liquid that may be on the mat surface. Then carefully remove the mat/seal. |
|  | - Wash the beads and elute the peptides. |
|  | - Wash the beads twice with 200 µL of 1 × PBS / 0.01% CHAPS and once with 200 µL of 0.1 × PBS / 0.01% CHAPS by mixing beads for 1.5 min in each wash solution. The peptides are eluted in PCR plates by mixing the beads for 5 min in 3% acetonitrile / 5% acetic acid / 50 mM citrate: |

The following lists plate components and positions on the KingFisher (note elution volume and buffer may change):

The following shows the details of the KingFisher method employed for washing and eluting the samples. No precollection was employed at the beginning of any step, the samples were not heated during mixing, and no postmixing was employed at the end of any step. The method takes about 25 minutes.

At conclusion of the Kingfisher method, there are three plates that require further processing (Capture plate #1, Elution plate #5, Bead collection plate #6)

|  | - Seal plate #1 (Capture plate) and store at 4°C. |
| --- | --- |
|  | - Seal plate #6 (Antibody-beads collection plate) and store at 4°C. Save collected beads. |

Processing of the Elution plate #5 (steps 11-13) is dependent on site.

Sites 1, 2

|  | - Cover the plate containing the eluates with ThermalSeal® PCR Sealing Film and spin the plate at 800 × g for 30 seconds. |
| --- | --- |
|  | - Place the elution plate on a plate magnet for 5 minutes. |
|  | - Transfer the eluates for LC-MS analysis. |
|  | - (Option 1) Split the eluates to two clean PCR plates by carefully drawing up 2 × 12 µL of the supernatant using a multi-channel pipet, without touching the bottom or sides of the wells, and transferring into corresponding wells of two clean PCR plates. Cover the plates with chemically resistant sealing foil, spin down the plates briefly, and store them at -80 °C. |
|  | - (Option 2) Transfer the eluates (20 µL each) to corresponding labeled autosampler vials and place in the LC autosampler (5° C) for LC-MRM analysis. |

Site 3

|  | - Desalt using Agilent RPS AssayMAP cartridges (refer to plate map below) |
| --- | --- |
|  | - Prime RP-S cartridges with 50 uL 90% ACN/0.1% TFA (10 uL/min). |
|  | - Equilibrate RP-S cartridges with 50 uL 0.1% TFA (10 uL/min). |
|  | - Load sample in dispense mode (2 uL/min). |
|  | - Wash cup with 25 uL 0.1% TFA. |
|  | - Wash sample with 50 uL 0.1% TFA (10 uL/min). |
|  | - Elute peptides with 50 uL 40% ACN/0.1% TFA (2 uL/min). |
|  | - (Manual) After protocol, transfer contents of each well of the elution plate to a corresponding labeled autosampler vial and freeze at -80 °C. |
|  | - (Manual) Vacuum centrifuge autosampler vials to dryness. Store at -80 °C until MS analysis. |

(below) Bravo AssayMAP plate map diagram for RPS desalting of immunoaffinity enriched peptides

| (1) Wash Station (connected to pumps primed with 0.1 Trifluoroacetic acid) | (2) AssayMAP Cartridges | (3) Waste Collection (96 Greiner 250 uL) |
| --- | --- | --- |
| (4) Sample [96 BioRad skirted] | (5) Priming (90% Acetonitrile/0.1% TFA) [12 column reservoir] | (6) Elution (40% Acetonitrile/0.1% TFA) [12 column reservoir] |
| (7) Flowthrough Collection [96 BioRad skirted] | (8) Equilibration (0.1% TFA) [12 column reservoir] | (9) Eluate Collection [96 BioRad skirted] |

Immunoaffinity enrichment using the second plex of antibodies

|  | - Retrieve plate #1 (Capture plate), see step 9 |
| --- | --- |
|  | - Remove seal cap/mat, and add 50 µL of the Antibody beads multi-mix #2 (mAbs to unmodified peptide targets) to each sample well. When adding the beads, mix the Antibody beads in the tube provided. Dispense enough beads to complete a capture plate into a separate tube. Vortex the bead solution and pipet the master mix bead solution up and down several times to ensure well-mixed beads. |
|  | - Proceed with overnight enrichment and sample processing as described in steps 4 and on. |

# Analysis of Samples by LC-MRM

## 9.1 Materials and Reagents

|  | LC autosampler | LC | MS | SS mix |
| --- | --- | --- | --- | --- |
| Site 1 | Dionex Ultimate 3000 RS autosampler | Dionex Ultimate 3000 RSLCnano | Quantiva | Pierce Retention Time Cal Mix |
| Site 2 | Eksigent 425 | Eksigent 425 | 5500 QTRAP | Pierce iRT |
| Site 3 | Proxeon Easy nLC-1000 | Proxeon Easy nLC-1000 | Quantiva | Michrom SSS |

|  | Site 1 | Site 2 | Site 3 |
| --- | --- | --- | --- |
| Mobile phase A | 2% ACN, 0.1% FA | 0.1% FA | 3% ACN / 0.1% FA |
| Mobile phase B | 90% ACN, 0.1% FA | 90% ACN/0.1%FA | 90% ACN / 0.1% FA |
| Trap Column | Acclaim PepmapTM 100, 100 µm × 2 cm, C18, 5 µm, 100A (Thermo) | ChromXP, C18, 5um, 0.3 × 5mm |  |
| Analytical Column | Acclaim PepmapTM RSLC, 75 µm × 25 cm, C18, 2 µm, 100A (Thermo) | ChromXP, C18, 3um, 0.075 × 150mm | 0.075 × 150 mm C18 Reprosil 1.9 um picofrit |
| Column Temperature | Trap at 30oC Analytical at RT | 45oC | 50oC |
| MS Resolution | Q1: 0.4, Q3: 0.7 | Unit/Unit | Q1: 0.4, Q3: 0.7 |
| MS Spray | 2200 V, Transfer tube 275° C | 3500 V | 2000V, ion transfer tube temp 300 |
| CID gas | 2 mTorr | Medium | 1.5 |
| Dwell time | Variable | Variable |  |
| RT window | Variable | 2 min | 8 min |
| Target cycle time |  | 1 sec | 1.5 sec |

Gradient:

| Site 1 |  |  | Site 2 |  |  |  | Site 3 |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Retention Time | Flow | %B | Retention Time | Flow | %B |  | Retention Time | Flow | %B |  |
| 0 | 0.3 | 2 | 0 | 5 | 2 |  | 0 | 0.2 | 0 |  |
| 0 | 0.3 | 2 | 3 | 5 | 2 |  | 2 | 0.2 | 6 |  |
| 7.5 | 0.3 | 2 | 3.1 | 0.3 | 2 |  | 47 | 0.2 | 30 |  |
| 8.5 | 0.3 | 8 | 4 | 0.3 | 2 |  | 53 | 0.2 | 60 |  |
| 50 | 0.3 | 26.5 | 5 | 0.3 | 10 |  | 54 | 0.2 | 90 |  |
| 56 | 0.3 | 50 | 25 | 0.3 | 35 |  | 59 | 0.2 | 90 |  |
| 56.5 | 0.3 | 90 | 28 | 0.3 | 90 |  |  |  |  |  |
| 62 | 0.3 | 96 | 30 | 0.3 | 90 |  |  |  |  |  |
| 63 | 0.3 | 2 | 31 | 0.3 | 2 |  |  |  |  |  |
| 78 | 0.3 | 2 | 40 | 0.3 | 2 |  |  |  |  |  |

## 9.2 Analysis Procedure

Site 1

|  | - Prepare MS System Suitability Standards and blank injections (3% ACN / 5% acetic acid / 50 mM citrate) in fresh autosampler vials. |
| --- | --- |
|  | - Inject 5 µL of 5 fmol/µL PRTC peptides using the gradient optimized for RAS iMRM targets. |
|  | - Set up scheduling using PRTC LC-MRM run and Skyline iRT library containing all RAS iMRM peptides. |
|  | - Analyze 6 µL of the sample. |

Site 2

|  | - Prepare MS System Suitability Standards and blank injections (3% ACN / 5% acetic acid / 50 mM citrate). |
| --- | --- |
|  | - Check system suitability (10 fmol on column Pierce iRT standard). |
|  | - Setup scheduling using RAS peptide mastermix. |
|  | - When ready for analysis, thaw the samples and spin the plate at 800 × g for 30 seconds. |
|  | - Add MS system suitability standards and blank injections (3% ACN / 5% acetic acid / 50 mM citrate) to empty wells for external system suitability checks. |
|  | - Cover the plate with a 96-well Silicone Sealing Mat, spin it down briefly, and place it on the autosampler. |
|  | - Analyze 10 µL of the sample. |

Site 3

|  | - Thaw autosampler vials for each plex. |
| --- | --- |
|  | - Add 10 µL of 3 % ACN / 5% acetic acid / 50 mM citrate. |
|  | - Vortex mix briefly, recap and centrifuge at 800 × g for 30s. |
|  | - Prepare autosampler vials for the blank (3% ACN / 5% acetic acid / 50 mM citric acid) and system suitability standard (stock solution 1 pmol/uL diluted to 50 fmol/uL with (3% ACN / 5% acetic acid / 50 mM citric acid. |
|  | - Analyze by scheduled MRM injecting 10 µL. |

# Data Analysis

|  | - Import the raw data into the Skyline document containing the targeted peptides. |
| --- | --- |
|  | - Integrate extracted ion chromatograms (XIC) of all transition ions for the peptide sequences in Skyline document. |
|  | - Check the integration and retention time alignment of the endogenous peptides with the heavy peptide standards and relative peak areas of transitions. Confirm identity of peptides by manual inspection of relative light/heavy fragment ion intensities and retention times. |
|  | - Export report data to .csv file. |
|  | - Sum Peak Area and Peak Background and sum transition ions for each peptide. |
|  | - Calculate Light to Heavy ratio of summed intensities. |

# Run Orders

The following is the run order for characterization plates.

Here is the order for running QC samples (Experiment 2, low, med, high):

| Day | Inject Order |
| --- | --- |
| 1 | L, M, H |
| 2 | H, L, M |
| 3 | M, H, L |
| 4 | L, H, M |
| 5 | M, L, H |

| Order | Sample Description | Location | Notes |
| --- | --- | --- | --- |
| 1 | SS (system suitability) |  | replace fields with values (e.g. "SSmix" with product name) |
| 2 | Plate 1 Low QC #1 | A1 | insert a date on the end if you like dates in your names |
| 3 | Plate 1 Med QC #1 | A2 |  |
| 4 | Plate 1 Hi Qc #1 | A3 |  |
| 5 | Plate 1 wash #1 |  |  |
| 6 | Plate 1 Reference Lysate #1 | A4 | insert any samples between these two reference runs |
| 7 | Plate 1 Reference Lysate #2 | H4 |  |
| 8 | Plate 1 wash #2 |  |  |
| 9 | Plate 1 Low QC #2 | E5 |  |
| 10 | Plate 1 Med QC #2 | F5 |  |
| 11 | Plate 1 Hi Qc #2 | G5 |  |
| 12 | Plate 1 wash #3 |  |  |
| 13 | Plate 1 Reference Lysate #3 | D5 |  |
| 14 | Plate 1 Reference Lysate #4 | H8 |  |
| 15 | Plate 1 wash #4 |  |  |
| 16 | Plate 1 Blank #1 | A9 |  |
| 17 | Plate 1 Blank #2 | B9 |  |
| 18 | Plate 1 Blank #3 | C9 |  |
| 19 | Plate 1 Curve B #1 | H10 |  |
| 20 | Plate 1 Curve C #1 | G10 |  |
| 21 | Plate 1 Curve D #1 | F10 |  |
| 22 | Plate 1 Curve E #1 | E10 |  |
| 23 | Plate 1 Curve F #1 | D10 |  |
| 24 | Plate 1 Curve G #1 | C10 |  |
| 25 | Plate 1 Curve H #1 | B10 |  |
| 26 | Plate 1 Curve I #1 | A10 |  |
| 27 | Plate 1 Carryover #1 |  |  |
| 28 | Plate 1 wash #5 |  |  |
| 29 | Plate 1 Low QC #3 | F9 |  |
| 30 | Plate 1 Med QC #3 | G9 |  |
| 31 | Plate 1 Hi Qc #3 | H9 |  |
| 32 | Plate 1 wash #6 |  |  |
| 33 | Plate 1 Reference Lysate #5 | D9 |  |
| 34 | Plate 1 Reference Lysate #6 | H12 |  |
| 35 | Plate 1 wash #7 |  |  |
| 36 | Plate 1 Beads only blank #1 | E1 |  |
| 37 | Plate 1 wash #8 |  |  |

# Fit-for-purpose validation

Table 1. Summary of performance figures of merit for analytical characterization of multiplexed assays in cell line lysates.

|  | LC- MRM | IMAC-MRM | Immuno-MRM |
| --- | --- | --- | --- |
| # analytes validated | 113 | 47 | 96 |
| Input amt (ug) | 1 | 200 | 500 |
| Min LLOQ (fmol/mg) | 80 | 2 | 0.82 |
| Max LLOQ(fmol/mg) | 8000 | 125 | 8000 |
| Median LLOQ (fmol/mg) | 200 | 12.5 | 6.12 |
| Median linear range (log10) | >3.7 | >2.9 | >3.4 |
| Median intra-assay %CV | 2.3 | 2.2 | 7.2 |
| Median inter-assay %CV | 2.7 | 2.3 | 15.9 |

Response curves were performed to characterize the linear range and determine limits of quantification. Heavy stable isotope-labeled peptides were serially diluted in a background matrix of pooled cell line lysates with a constant concentration of light synthetic peptides. Peptides were analyzed with or without antibody or immobilized metal bead enrichment and the peak area ratio (heavy : light) was plotted as a function of spiked peptide concentration. Validated assays were required to show a linear response in the concentrations measured and lower limits of quantification (LLOQ) were determined by the lowest point in the curve measured with a CV less than 20%. The linear range observed for each set of assays was greater than three orders of magnitude.

The expected precision of the assays was characterized by performing repeatability experiments to measure intra-assay (within day) and inter-assay (between day) variability. Three concentrations of heavy stable isotope-labeled peptides (low, medium, high) were spiked into aliquots of a pooled background of cell line lysates. Light synthetic peptides were added at a constant concentration in each aliquot. Complete process replicates (including digestion) were measured using each multiplexed assay over five days. For inclusion in precision characterization, peptides were required to be above the LLOQ in 2/3 replicates on 3/5 days. A summary of the repeatability of the assays is shown in Table 1. As expected, inter-assay variation is slightly higher compared to intra-assay variation. The Immuno-MRM assays target the lowest abundance peptides and require an extra step of sample processing, likely leading to slightly higher reported variation.

The success rates for assay development were 113/116 (97%) for LC-MRM, 47/49 (96%) in IMAC-MRM, and 96/103 (93%) in Immuno-MRM. Peptides to fail in validation were predominantly caused by low signal intensity or interference leading to high imprecision. This could be due to poor ionization efficiency in the mass spectrometer, poor recovery of the peptide during enrichment, or poor affinity of the antibodies. The three peptides that failed in LC-MRM were SQQLSNDHICYFLYQILR (MAPK3), ALDLLDR (MAPK3), and DSSYYWEIEASEVMLSTR (RAF1). The two phosphopeptides that failed in IMAC-MRM were doubly phosphorylated, TAGTSFMMpTPpYVVTR (MAPK8) and HTDDEMpTGpYVATR (MAPK14). The seven peptides that failed in immuno-MRM were largely related to phosphorylation: the doubly phosphorylated YpSSDPTGALTEDpSIDDTFLPVPEYINQSVPK (EGFR), LLDIDETEpYHADGGK (ERBB2), LLDIDETEYHADGGK (ERBB2), TAGTSFMMpTPpYVVTR (MAPK14), TAGTSFMMpTPYVVTR (MAPK14), AVpSMDNSNK (FOXO3), DEVAHTLTESR (AKT3). Previous characterization of the antibodies (performed during antibody development) indicated the ERBB2 (LLDI, CPTAC-ERBB-3) and AKT3 (DEVA, CPTAC-AKT3-7) antibodies should recognize the unmodified peptide sequences. ERBB2 (LLDI peptide) was only detected in 1 out of 3 sites using the immuno-MRM assay and the AKT3 (DEVA peptide) failed in the repeatability experiment due to low signal intensity.