# STANDARD OPERATING PROCEDURE

Title: Peptide immunoaffinity enrichment for the characterization of intra- and inter-day repeatability of immuno-MRM assays (Validation Samples). Version #: 2018; Author: Paulovich Lab.

# Purpose

The purpose of this document is to describe the immunoaffinity enrichment of peptides as part of the characterization of the intra-day and inter-day repeatability of immuno-MRM assays over 5 days.

# Scope

This procedure may be used to generate immunoaffinity-enriched peptide repeatability (validation) samples to be analyzed by MRM.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- KingFisher 96 magnetic bead handling platform (Thermo Scientific), equipped with a PCR magnet head

- LabQuake tube rotator (415110, Barnstead)

- Sepharose Magnetic Protein G beads (28-9513-79, GE)

- DynaMagTM-96 Side Skirted Magnet (12027, Life Technologies)

- Centrifuge for use with 96 well plates (5810R, Eppendorf)

# Materials

- 1.7 mL microcentrifuge tubes (24-282, Genesee Scientific)

- 15 mL conical centrifuge tubes (352196, Corning Life Sciences)

- MColorpHast® pH Test Strips, 5-10 (1.09533.0001, EMD Millipore)

- KingFisher 96 KF microplate (200 L) (97002540, Thermo Scientific)

- KingFisher 96 tip comb for PCR magnets (97002514, Thermo Scientific)

- Hard-Shell® Low-Profile Thin-Wall 96-Well Skirted PCR Plates (HSP-9601, BioRad)

- Chemically resistant sealing foil (T-3025-8B, BioExpress)

- ThermalSeal® PCR Sealing Films (12-168, Genesee Scientific)

- 1 M Trizma hydrochloride buffer, pH 8.0 (Tris, T2694, Sigma)

- 10X Phosphate buffered saline (PBS) (BP39920, Fisher Scientific)

- CHAPS detergent (28300, Thermo Scientific)

- Water, Optima® LC/MS, suitable for UHPLC-UV (W6-4, Fisher Scientific)

- Acetic acid (ACS reagent, ≥99.7%, 242853, Sigma)

- Acetonitrile, Optima® LC/MS, suitable for UHPLC-UV (A955-4, Fisher Scientific)

- Formic acid (FA) (1.11670.1000, EMD Millipore)

- Sodium azide (NaN3) (S8032, Sigma)

# Reagents

- Light and stable isotope-labeled heavy synthetic peptides for MRM analyses were purchased at >95% pure as measured by HPLC (“pure”) or ~75% purity (“crude”). Cysteine residues of the tryptic light and heavy peptides were carbamidomethylated, and the heavy peptides’ C-terminal lysine (K) or arginine (R) residues were uniformly 13C and 15N labeled, yielding a mass increase of 8 Da (for K) or 10 Da (for R). For peptides at the C-terminus of proteins, other amino acids (such as valine) can be isotopically labeled. Amino acid analysis (AAA) was performed to determine peptide concentrations. Individual peptide stock aliquots were stored in 30% acetonitrile-0.1% formic acid at 100 pmol/L at -80 °C.

- Rabbit-anti-human and mouse-anti-human monoclonal antibodies (mAbs) were custom-generated to the tryptic and proteotypic peptide targets and covalently coupled to magnetic Protein G beads.

- Undigested background lysate from a pooled mix of tumor tissue lysates.

# Solutions

- Prepare a 5% (w/v) CHAPS solution in Optima® water by dissolving 1.0 g of CHAPS powder in 20 mL of Optima® water. It will take some time to dissolve the powder completely. (CHAPS aids in the beads transfer from plate to plate on the KingFisher.)

- Prepare 100 mL of a 1X PBS-0.01% CHAPS solution, pH 7 by adding 10.0 mL of 10X PBS and 200 L of 5% CHAPS to Optima® water, for a total of 100 mL.

- Prepare 100 mL of a 1/10X PBS-0.01% CHAPS solution, pH 7 by adding 1.0 mL of 10X PBS and 200 L of 5% CHAPS to Optima® water, for a total of 100 mL.

- Prepare 100 mL of 3% acetonitrile-5% acetic acid by adding 3.0 mL of Optima® acetonitrile and 5.0 mL of glacial acetic acid to Optima® water, for a total of 100 mL.

- Prepare a 10% (w/v) NaN3 in water by dissolving e.g. 1.0 g of NaN3 powder in 10 mL of Optima® water.

- Prepare 50 mL of a 1X PBS-0.03% CHAPS-0.1% NaN3 solution by adding 5.0 mL of 10X PBS, 300 L of 5% CHAPS, and 500 L of 10% NaN3 to Optima® water, for a total of 50 mL.

- Prepare 1 L of 3% acetonitrile-0.1% formic acid by adding 30.0 mL of Optima® acetonitrile and 1.0 mL of neat formic acid to Optima® water, for a total of 1 L.

# Procedure

The following describes experiments for the intra- and inter-day repeatabilities of a multiplexed set of immuno-MRM assays whose lower limits of quantification (LLOQs) and linear ranges have previously been determined using response curves. The repeatability experiments are performed on 5 different days, and each assay is tested at three different analyte levels (low, med, high) per day, with three replicates performed at each analyte level per day (i.e., for a total of 45 experiments). See also the CPTAC assay portal’s guidance document for more information (https://assays.cancer.gov/guidance-document/).

As was done for the response curve experiments, the stable isotope-labeled heavy synthetic peptides are used as analytes in these experiments, and the light peptides are used as internal standards, since the endogenous detection of some of the peptides in matrix does not make it possible to use the light peptides as the analyte peptides.

Preparation of the repeatability samples for one day’s experiments:

Briefly, the following considerations were taken into account when preparing the samples:

- the volume of peptides + matrix sample per capture is 200 L, and

- it is convenient to perform the dilutions from the hi samples to the lower level samples in bulk first, and only after the dilutions have been made are the samples aliquoted to arrive at 3 replicate samples per concentration level.

- A heavy peptide master mix was aliquoted into at least 5 microcentrifuge tubes and stored at -80 °C, with one tube to be thawed for a day’s experiments.

- Thaw a tube containing background matrix and digest enough lysate to prepare 3x replicates of low, medium, and high concentrations.

- Spike light peptide with a concentration of 80 fmol peptide per sample,

- Add heavy peptides, and perform the dilutions as follows:

  - Hi 800 fmol

  - Medium 80 fmol

  - Low 8 fmol

- Split each concentration level into three replicates.

Perform immunoaffinity enrichment (capture) and LC-MRM of the peptides from the repeatability samples according to SOPs.

Sample run order on the LC-MRM-MS system:  The repeatability samples are run in increasing order of heavy peptide level, with one injection per capture replicate. Before that, a system suitability standard peptide mix is run to ascertain that the system is behaving as expected (see also Abbatiello, S. E., et al.1). An injection of the synthetic light peptide standards in buffer is then performed to determine the approximate retention times of the peptides, with which a scheduled LC-MRM method is generated using the Skyline software2. This is followed by the samples, with a buffer blank run (injecting 10 L of just 3% acetonitrile-5% acetic acid) in between each sample run. (Due to some peptides showing considerable carry-over from run to run, it is not practical to randomize the run order with respect to heavy peptide spike level.)

# Referenced Documents

1.	Abbatiello, S. E., Mani, D. R., Schilling, B., Maclean, B., Zimmerman, L. J., Feng, X., Cusack, M. P., Sedransk, N., Hall, S. C., Addona, T., Allen, S., Dodder, N. G., Ghosh, M., Held, J. M., Hedrick, V., Inerowicz, H. D., Jackson, A., Keshishian, H., Kim, J. W., Lyssand, J. S., Riley, C. P., Rudnick, P., Sadowski, P., Shaddox, K., Smith, D., Tomazela, D., Wahlander, A., Waldemarson, S., Whitwell, C. A., You, J., Zhang, S., Kinsinger, C. R., Mesri, M., Rodriguez, H., Borchers, C. H., Buck, C., Fisher, S. J., Gibson, B. W., Liebler, D., MacCoss M., Neubert, T. A., Paulovich, A., Regnier, F., Skates, S. J., Tempst, P., Wang, M., and Carr, S. A. (2013) Design, implementation and multisite evaluation of a system suitability protocol for the quantitative assessment of instrument performance in liquid chromatography-multiple reaction monitoring-MS (LC-MRM-MS). Mol Cell Proteomics 12, 2623-2639.

2.	MacLean, B., Tomazela, D. M., Shulman, N., Chambers, M., Finney, G. L., Frewen, B., Kern, R., Tabb, D. L., Liebler, D. C., and MacCoss, M. J. (2010) Skyline: an open source document editor for creating and analyzing targeted proteomics experiments. Bioinformatics 26, 966-968.