# STANDARD OPERATING PROCEDURE

Title: Peptide immunoaffinity enrichment for the characterization of immuno-MRM assays using response curves. Version #: 2018; Author: Paulovich Lab.

# Purpose

The purpose of this document is to describe generation of response curves for characterization of linearity and lower limits of quantification (LLOQ).

# Scope

This procedure may be used to generate immunoaffinity-enriched peptide response curve samples to be analyzed by MRM. It includes a description of generating and analyzing the response curve samples.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- KingFisher 96 magnetic bead handling platform (5400500, Thermo Scientific), equipped with a PCR magnet head

- LabQuake tube rotator (415110, Barnstead)

- Dynabeads® MPC®-S (Magnetic Particle Concentrator, A13346, Life Technologies) or DynaMagTM-15 Magnet (12301D, Life Technologies)

- DynaMagTM-96 Side Skirted Magnet (12027, Life Technologies)

- Centrifuge for use with 96 well plates (5810R, Eppendorf)

# Materials

- 1.7 mL microcentrifuge tubes (24-282, Genesee Scientific)

- 15 mL conical centrifuge tubes (352196, Corning Life Sciences)

- MColorpHast® pH Test Strips, 5-10 (1.09533.0001, EMD Millipore)

- KingFisher 96 KF microplate (200 L) (97002540, Thermo Scientific)

- KingFisher 96 tip comb for PCR magnets (97002514, Thermo Scientific)

- Hard-Shell® Low-Profile Thin-Wall 96-Well Skirted PCR Plates (HSP-9601, BioRad)

- Chemically resistant sealing foil (T-3025-8B, BioExpress)

- ThermalSeal® PCR Sealing Films (12-168, Genesee Scientific)

- 1 M Trizma hydrochloride buffer, pH 8.0 (Tris, T2694, Sigma)

- 10X Phosphate buffered saline (PBS) (BP39920, Fisher Scientific)

- CHAPS detergent (28300, Thermo Scientific)

- Water, Optima® LC/MS, suitable for UHPLC-UV (W6-4, Fisher Scientific)

- Acetic acid, ACS reagent, ≥99.7% (242853, Sigma)

- Acetonitrile, Optima® LC/MS, suitable for UHPLC-UV (A955-4, Fisher Scientific)

- Sodium azide (NaN3) (S8032, Sigma)

- Lima bean trypsin inhibitor (LS002829, Worthington Biochemical Corporation)

- Bovine 6 protein tryptic digest equal molar mix (PTD/00006/63, Bruker)

# Reagents

- Light and stable isotope-labeled heavy synthetic peptides for MRM analyses were purchased at >95% purity as measured by HPLC. Cysteine residues of the tryptic light and heavy peptides were carbamidomethylated, and the heavy peptides’ C-terminal lysine (K) or arginine (R) residues were uniformly 13C and 15N labeled, yielding a mass increase of 8 Da (for K) or 10 Da (for R). For peptides at the C-terminus of proteins, other amino acids (such as valine) can be isotopically labeled. Amino acid analysis (AAA) was performed to determine peptide concentrations. Individual peptide stock aliquots were stored in 30% acetonitrile-0.1% formic acid at 100 pmol/L at -80 °C.

- Rabbit-anti-human monoclonal antibodies (mAbs) were custom-generated to the tryptic and proteotypic peptide targets and covalently coupled to magnetic Protein G beads.

- Background matrix (e.g. lysate or plasma, as appropriate for sample set)

# Solutions

- Prepare a 5% (w/v) CHAPS solution in Optima® water by dissolving 1.0 g of CHAPS powder in 20 mL of Optima® water. It will take some time to dissolve the powder completely. (CHAPS aids in the bead transfer from plate to plate on the KingFisher.)

- Prepare 100 mL of a 1X PBS-0.01% CHAPS solution, pH 7 by adding 10.0 mL of 10X PBS and 200 L of 5% CHAPS to Optima® water, for a total of 100 mL.

- Prepare 100 mL of a 1/10X PBS-0.01% CHAPS solution, pH 7 by adding 1.0 mL of 10X PBS and 200 L of 5% CHAPS to Optima® water, for a total of 100 mL.

- Prepare 100 mL of 3% acetonitrile-5% acetic acid by adding 3.0 mL of Optima® acetonitrile and 5.0 mL of glacial acetic acid to Optima® water, for a total of 100 mL.

- Prepare a 10% (w/v) NaN3 solution in water by dissolving e.g. 1.0 g of NaN3 powder in 1.0 mL of Optima® water.

- Prepare 50 mL of a 1X PBS-0.03% CHAPS-0.1% NaN3 solution by adding 5.0 mL of 10X PBS, 300 L of 5% CHAPS, and 500 L of 10% NaN3 to Optima® water, for a total of 50 mL.

# Procedure

The following describes the experiments for a response curve that has varying concentrations of the heavy (analyte) peptides and a constant concentration of the light (internal standard) peptides in a human cell line matrix, including a 0 fmol (blank) sample that does not contain heavy peptides. The light peptides were not chosen as analytes since the endogenous detection of some of the peptides in background matrix would not have made it possible to determine the limits of detection and quantification.

Preparation of the response curve samples

This procedure describes heavy peptides that were tested in these particular response curves at 5000, 1000, 200, 40, 8, 2, 0.5, 0.125, and 0 fmol per 500 g of lysate digest. These levels were generated using serial dilutions, with the dilutions from 5000 to 8 fmol/10 L plasma being covered by 1:5 serial dilutions, and from 8 to 0.125 fmol/10 L plasma by 1:4 serial dilutions (smaller intervals were chosen for the lower levels since this would allow more robust characterizations of the limits of quantification). The synthetic light peptides were spiked at a constant level of 200 fmol per sample.

- Thaw a tube containing background digest lysate.

- Prepare a light peptide-digest mix as follows (this yields a total volume of 8600 L):

7.1. 	Add 7738 L of 1X PBS-0.01% CHAPS to a 15 mL conical centrifuge tube;

7.2. 	add 430 L of digest;

7.3.	add 430 L of 1 M Tris, pH 8.0 (to ensure a pH between 7.0 and 8.0);

7.4.	add 2.15 L of the 2 pmol/L synthetic light peptide master mix (yielding a synthetic light peptide concentration of 200 fmol/sample).

7.5.	vortex and spin down the sample briefly.

- Aliquot the light peptide-digest mix, add heavy peptides, and perform the serial dilutions as follows:

8.1.	Aliquot the light peptide-digest mix into eight 1.7 mL microcentrifuge tubes:

a) 1100 L into 1 tube (label the tube with 5000 fmol)

b) 880 L into each of 4 tubes (label the tubes with 1000, 200, 40, and 8 (fmol))

c) 660 L into each of 3 tubes (label the tubes with 2, 0.5, and 0.125 (fmol)).

8.2.	Place the tube with the remaining light peptide-digest mix on ice (5 x 200 L of this mix will later be transferred to 5 wells of a 96 well plate, representing the 0 fmol replicate samples).

8.3.	Take 13.75 L out of the first (5000) tube and discard, and add 13.75 L of the 2 pmol/L heavy peptide master mix to that tube; mix by centrifuging and spin down briefly. Since the heavy peptide mix contained formic acid, measure the pH of the peptide+plasma sample using e.g. 2 L spotted onto a pH 5.0-10.0 pH test strip to ensure a pH above 7.0. The heavy peptide concentration in this sample is 5000 fmol.

8.4.	Perform the 1:5 and 1:4 serial dilutions using 220 L from the first tube and diluting first in the tubes containing 880 L of the light peptide-digest mixes, and then into the tubes containing 660 L. Ensure good mixing at each dilution step.

Perform immunoaffinity enrichment (capture) of the peptides from the digest samples according to the enrichment SOP and analyze by LC-MRM.

Sample run order on the LC-MRM-MS system: The response curve samples are run in increasing order of heavy peptide level, with one injection per capture replicate. Before that, a system suitability standard peptide mix is run to ascertain that the system is behaving as expected (see also Abbatiello, S. E., et al.1). An injection of the synthetic light peptide standards in buffer is then performed to determine the approximate retention times of the peptides, with which a scheduled LC-MRM method is generated using the Skyline software2. This is followed by the samples, with a buffer blank run (injecting 10 L of 3% acetonitrile-5% acetic acid) in between each sample run. (Due to some peptides showing considerable carry-over (e.g. >1%) from run to run, it is not practical to randomize the run order with respect to heavy peptide spike level.)

# Referenced Documents

1.	Abbatiello, S. E., Mani, D. R., Schilling, B., Maclean, B., Zimmerman, L. J., Feng, X., Cusack, M. P., Sedransk, N., Hall, S. C., Addona, T., Allen, S., Dodder, N. G., Ghosh, M., Held, J. M., Hedrick, V., Inerowicz, H. D., Jackson, A., Keshishian, H., Kim, J. W., Lyssand, J. S., Riley, C. P., Rudnick, P., Sadowski, P., Shaddox, K., Smith, D., Tomazela, D., Wahlander, A., Waldemarson, S., Whitwell, C. A., You, J., Zhang, S., Kinsinger, C. R., Mesri, M., Rodriguez, H., Borchers, C. H., Buck, C., Fisher, S. J., Gibson, B. W., Liebler, D., MacCoss, M., Neubert, T. A., Paulovich, A., Regnier, F., Skates, S. J., Tempst, P., Wang, M., and Carr, S. A. (2013) Design, implementation and multisite evaluation of a system suitability protocol for the quantitative assessment of instrument performance in liquid chromatography-multiple reaction monitoring-MS (LC-MRM-MS). Mol Cell Proteomics 12, 2623-2639.

2.	MacLean, B., Tomazela, D. M., Shulman, N., Chambers, M., Finney, G. L., Frewen, B., Kern, R., Tabb, D. L., Liebler, D. C., and MacCoss, M. J. (2010) Skyline: an open source document editor for creating and analyzing targeted proteomics experiments. Bioinformatics 26, 966-968.