# STANDARD OPERATING PROCEDURE

Title: Peptide immunoaffinity enrichment for the characterization of immuno-MRM assays using response curves. Version #: 2021; Author: Paulovich Lab.

# Purpose

The purpose of this document is to describe generation of response curves for characterization of linearity and lower limits of quantification (LLOQ).

# Scope

This procedure may be used to generate immunoaffinity-enriched peptide response curve samples to be analyzed by MRM.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- KingFisher 96 magnetic bead handling platform (5400500, Thermo Scientific), equipped with a PCR magnet head

- Titer plate shaker, Lab-line Instrument

- Protein G Mag Sepharose beads Xtra (28967070, Cytiva)

- Centrifuge for use with 96 well plates (5810R, Eppendorf)

# Materials

- MColorpHast® pH Test Strips, 5-10 (1.09533.0001, EMD Millipore)

- KingFisher 96 KF microplate (200 L) (97002540, Thermo Scientific)

- KingFisher 96 tip comb for PCR magnets (97002514, Thermo Scientific)

- Hard-Shell® Low-Profile Thin-Wall 96-Well Skirted PCR Plates (HSP-9601, BioRad)

- Non-stick square film. (961201, Analytic)

- ThermalSeal® PCR Sealing Films (12-168, Genesee Scientific)

- 1 M Trizma hydrochloride buffer, pH 8.0 (Tris, T2694, Sigma)

- 10X Phosphate buffered saline (PBS) (BP39920, Fisher Scientific)

- CHAPS detergent (28300, Thermo Scientific)

- Water, Optima® LC/MS, suitable for UHPLC-UV (W6-4, Fisher Scientific)

- Acetic acid, ACS reagent, ≥99.7% (242853, Sigma)

- Acetonitrile, Optima® LC/MS, suitable for UHPLC-UV (A955-4, Fisher Scientific

# Reagents

- Light peptides were purchased at crude purity (~85%). Cleavable stable isotope-labeled heavy synthetic peptides for MRM analyses were purchased at >95% purity by HPLC. Cysteine residues of the tryptic light and heavy peptides were carbamidomethylated, and the heavy peptides’ tryptic C-terminal lysine (K) or arginine (R) residues were uniformly 13C and 15N labeled, yielding a mass increase of 8 Da (for K) or 10 Da (for R). For peptides at the C-terminus of proteins, other amino acids (such as valine) are isotopically labeled. Individual peptide stock aliquots were stored in 30% acetonitrile-0.1% formic acid at 100 pmol/L at -80 °C.

- Rabbit-anti-human monoclonal antibodies (mAbs) were custom-generated to the tryptic and proteotypic peptide targets and covalently coupled to magnetic Protein G beads.

- Background matrix (e.g., pooled tissue lysate or plasma, as appropriate for sample set)

# Solutions

- Prepare a 5% (w/v) CHAPS solution in Optima® water by dissolving 1.0 g of CHAPS powder in 20 mL of Optima® water.

- Prepare 100 mL of a 1X PBS-0.01% CHAPS solution, pH 7 by adding 10.0 mL of 10X PBS and 200 L of 5% CHAPS to Optima® water, for a total of 100 mL.

- Prepare 100 mL of a 1/10X PBS-0.01% CHAPS solution, pH 7 by adding 1.0 mL of 10X PBS and 200 L of 5% CHAPS to Optima® water, for a total of 100 mL.

- Prepare 100 mL of 3% acetonitrile-5% acetic acid by adding 3.0 mL of Optima® acetonitrile and 5.0 mL of glacial acetic acid to Optima® water, for a total of 100 mL.

# Procedure

|  | 1; 2; 3; 4; 5; 6; 7. Per sample (fmol); Total samples; Vol in well; H peptide(fmol/ul); Stock (uL); Vol in the tube after serial dilution; Conc. (fmol/ul). A; 2000.00; 5; 64.00; 1000.00; 15.20; 70.80; 200.000. B; 200.00; 5; 69.96; 8.40; 66.40; 20.000. C; 20.00; 5; 99.59; 11.95; 64.03; 2.000. D; 8.00; 5; 98.98; 47.51; 99.71; 0.800. E; 3.20; 5; 97.44; 46.77; 99.28; 0.320. F; 1.28; 5; 93.60; 44.93; 98.21; 0.128. G; 0.51; 5; 84.00; 40.32; 95.52; 0.051. H; 0.20; 5; 60.00; 28.80; 88.80; 0.020. |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Table 1 |  |  |  |  |  |  |  |

- Thaw a tube containing Heavy peptides and make 1 pmol/ul stock solution.

- Thaw a tube containing Light peptides and make 1 pmol/ul stock solution.

- Thaw a tube containing cell/tissue lysate and plate on deep 96 well plate, amount can vary depended on experiment, e,g, 500 ug or 100 ug per well.

4.     Follow the Table 1 to make heavy peptide standard curve for 5 replicates.

4.1. 	    Plate 0.03% diluted yeast lysate in well according Column 3 in Table 1.

4.2. 	    Add 15.2 heavy peptide solution to A well, Table 1.

4.3.	    Serial transfer liquid from well A to H following Column 6, Table 1.

4.4.	    For 5 blank samples, only add Light peptides.

4.5.        Make a 20 fmol/ul Light peptide mix enough to do 45 immuno captures.

5.     Perform trypsin Digestion on curve samples, see SOP of trypsin digestion.

6.     Add 10 ul of Light peptide to curve samples prior to desalting.

7.     Add antibody to the curve samples, see SOP on immunocapture.

8.     LC-MRM analysis of curve samples, see SOP on LC-MRM.

1.	Abbatiello, S. E., Mani, D. R., Schilling, B., Maclean, B., Zimmerman, L. J., Feng, X., Cusack, M. P., Sedransk, N., Hall, S. C., Addona, T., Allen, S., Dodder, N. G., Ghosh, M., Held, J. M., Hedrick, V., Inerowicz, H. D., Jackson, A., Keshishian, H., Kim, J. W., Lyssand, J. S., Riley, C. P., Rudnick, P., Sadowski, P., Shaddox, K., Smith, D., Tomazela, D., Wahlander, A., Waldemarson, S., Whitwell, C. A., You, J., Zhang, S., Kinsinger, C. R., Mesri, M., Rodriguez, H., Borchers, C. H., Buck, C., Fisher, S. J., Gibson, B. W., Liebler, D., MacCoss, M., Neubert, T. A., Paulovich, A., Regnier, F., Skates, S. J., Tempst, P., Wang, M., and Carr, S. A. (2013) Design, implementation and multisite evaluation of a system suitability protocol for the quantitative assessment of instrument performance in liquid chromatography-multiple reaction monitoring-MS (LC-MRM-MS). Mol Cell Proteomics 12, 2623-2639.

2.	MacLean, B., Tomazela, D. M., Shulman, N., Chambers, M., Finney, G. L., Frewen, B., Kern, R., Tabb, D. L., Liebler, D. C., and MacCoss, M. J. (2010) Skyline: an open source document editor for creating and analyzing targeted proteomics experiments. Bioinformatics 26, 966-968.