# STANDARD OPERATING PROCEDURE

Title: Liquid Chromatography – Multiple Reaction Monitoring for targeted analysis of peptides. Version #: 2017; Author: Paulovich Lab.

# Purpose

The purpose of this document is to describe the configuration of the LC-MRM system for targeted proteomics analysis.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Eksigent 425 LC system, incl 425 autosampler

- 5500 QTRAP mass spectrometer

# Conditions

|  | Description |
| --- | --- |
| Mobile phase A | 0.1% FA |
| Mobile phase B | 90% ACN/0.1%FA |
| Trap Column | Reprosil, C18, 3um, 0.3 × 5mm |
| Analytical Column | ChromXP, C18, 3um, 0.075 × 150mm |
| Column Temperature | 45oC |
| MS Resolution | Unit/Unit |
| MS Spray | 1300 V |
| CID gas | Medium |
| Dwell time | Variable |
| RT window | 240 sec |
| Target cycle time | 1.5 sec |

Gradient:

| Fred Hutch |  |  |
| --- | --- | --- |
| Retention Time | Flow | %B |
| 0 | 5 | 1 |
| 4 | 5 | 1 |
| 6 | 0.3 | 6 |
| 31 | 0.3 | 25 |
| 36 | 0.3 | 35 |
| 38 | 0.3 | 90 |
| 40 | 0.3 | 90 |
| 42 | 0.3 | 1 |
| 50 | 0.3 | 1 |

# Procedure

- Prepare MS System Suitability Standards and blank injections (3% ACN / 5% acetic acid / 50 mM citrate).

- Check system suitability (5 fmol on column Pierce iRT standard).

- Setup scheduling using synthetic peptide mastermix.

- When ready for analysis, thaw the samples and spin the plate at 800 × g for 30 seconds.

- Add MS system suitability standards and blank injections (3% ACN / 5% acetic acid / 50 mM citrate) to empty wells for external system suitability checks.

- Cover the plate with a 96-well Silicone Sealing Mat, spin it down briefly, and place it on the autosampler.

- Analyze 10 µL of the sample.

- Import the raw data into the Skyline document containing the targeted peptides.

- Integrate extracted ion chromatograms (XIC) of all transition ions for the peptide sequences in Skyline document.

- Check the integration and retention time alignment of the endogenous peptides with the heavy peptide standards and relative peak areas of transitions. Confirm identity of peptides by manual inspection of relative light/heavy fragment ion intensities and retention times.

- Upload final skyline files to Panorama.