# STANDARD OPERATING PROCEDURE

Title: Liquid Chromatography and PRM mass spectrometry, Easy-nLC II and Q Exactive Orbitrap. Version #: 1; Author: Yiying Zhu. Date: 2023/12.

# Purpose

The purpose of this document is to describe the liquid chromatography (LC) method and the mass spectrometry (MS) method for quantitative analysis of peptides using parallel reaction monitoring (PRM)

# Scope

This procedure may be used to set LC and MS method parameters for PRM target analysis of peptides.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- HPLC: EASY-nLC II (Thermo Scientific)

- Mass spectrometer: Q Exactive Orbitrap (Thermo Scientific)

# Materials

- Injection loop 20 µL (Thermo Scientific)

- Ionization Source Picoview® (New Objective)

- Analytical column (ID 100 µm, 3 µm, 15 cm), self-packed

- Water, Burdick &Jackson® HPLC grade (Fisher)

- Acetonitrile (ACN), OmniSolv® Gradient grade (Fisher)

- Formic acid (FA), Optima® LC-MS grade (Fisher)

- Trifluoroacetic acid (TFA)

# Reagents

- Standards: Heavy and light stable isotope-labeled standard peptides (SIS) are synthesized in the peptide lab of Cell Signaling Technology. All peptides are purified by HPLC, and amino acid analysis (AAA) is performed to determine peptide concentrations. Cysteine residues of the peptides are carbamidomethylated. All heavy-SIS are labeled at internal amino acids. The isotope-labeled amino acid positions are shown in the following table and marked in red. Phosphorylated amino acids are followed by a (p).

Heavy-SIS

RPHFPQFS(p)YSASGTA

IGEGTY(p)GVVYK

LNT(p)SDFQK

YCRPESQEHPEADPGSAAPY(p)LK

AVDGY(p)VKPQIK

GY(p)VPATIK

IADPEHDHTGFLT(p)EY(p)VATR

VADPDHDHTGFLT(p)EY(p)VATR

VADPDHDHTGFLT(p)EYVATR

LAS(p)PELER

# Procedure

- Set up LC method parameters

  - Sample pickup: Volume 4 µL; Flow 10 µL /min

  - Sample loading: Volume 7 µL; Max pressure 280 bar

  - Gradient

| Time | Duration | Flow (nL/min) | %B |
| --- | --- | --- | --- |
| 00:00 | 00:00 | 300 | 0 |
| 01:00 | 01:00 | 300 | 9 |
| 31:00 | 30:00 | 300 | 32 |
| 34:00 | 03:00 | 300 | 98 |
| 40:00 | 06:00 | 300 | 98 |

  - Analytical column equilibration

    - Volume 9 µL

    - Max pressure 280.00 Bar

  - Autosampler wash: Standard flush volume 100.00 µL of solvent A

- Set up tune file parameters and MS method parameters

- NSI source parameters

    - Sheath gas flow rate 0

    - Aux gas flow rate 0

    - Sweep gas flow rate 0

    - Spray voltage kv 2.1

    - Capillary temp. ˚C 275

    - S-lens RF level 55.0

- MS parameters

  - Micro scans 1

  - Resolution 17500

  - AGC target 5e5

  - Maximum IT 120 ms

  - Isolation window 1.5 m/z

  - Lock mass 371.10123

- Test system suitability with a peptide standard mix

Load 100 fmol of Bovine Serum Albumin (BSA) tryptic digest on the column and run it using the PRM method. A few BSA peptides are monitored, and their peaks are integrated and assessed in Skyline software to ensure the reproducibility and sensitivity of the system based on historical data. https://skyline.ms/project/home/software/skyline/begin.view If peak shape or peak intensities are unacceptable, troubleshoot, change the column, or clean the MS.

- Identify scheduling times for target peptides

- Prior to analysis, data-dependent analysis of light-SIS (sample loading amount: 100 fmol of each peptide) is performed on Q Exactive Orbitrap to generate a spectral library. The raw file is also imported into Skyline software and used to select the best charge state and PRM transitions for each peptide.

- Timing the peptide detection

  - Inject light-SIS mix (1 fmol of each peptide) to a mass spectrometer using an unscheduled PRM acquisition method

  - Import the raw data file into Skyline

  - Check the integration of all peaks and adjust manually if needed

  - Export a scheduled isolation list from Skyline with a 7-min window for each peptide

  - Directly import the isolation table to the inclusion list of the PRM method

  - Run light-SIS (1 fmol) again using the scheduled PRM acquisition method

  - Import the acquired data into Skyline to recheck retention times