# STANDARD OPERATING PROCEDURE

Title: Peptide immuno-affinity enrichment for the characterization of repeatability of immuno-PRM assays (Validation samples preparation). Version #: 1; Author: Yiying Zhu. Date: 2023/12.

# Purpose

The purpose of this document is to describe the characterization of repeatability of multiplex immuno-PRM assays. This is to estimate the performance of the assays over 5 days

# Scope

This procedure may be used to cover the overall preparation and running of samples to evaluate the repeatability/validation of immuno-affinity enrichment and peptide analysis by Parallel Reaction Monitoring in a complex sample.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Microcentrifuge

- Vacuum centrifuge

- HPLC: EASY-nLC II (Thermo Scientific)

- Mass spectrometer: Q exactive Orbitrap (Thermo Scientific)

# Materials

- Water, Burdick &Jackson® HPLC grade (Fisher)

- Acetonitrile, OmniSolv® Gradient grade (Fisher)

- Formic acid, Optima® LC-MS grade (Fisher)

- Trifluoroacetic acid (TFA)

- Protein A agarose beads slurry 50 % (Cell Signaling Technology, #9863)

- Stage tip/tube assembly

Reagents

- Standards: Heavy and light isotope-labeled standard peptides (SISs) are synthesized in the peptide lab at Cell Signaling Technology (CST). All peptides are purified by HPLC, and amino acid analysis (AAA) is performed to determine peptide concentrations. Cysteine residues of the peptides are carbamidomethylated. All heavy-SIS are labeled at internal amino acids. The isotope-labeled amino acid positions are shown in “SOP_Liquid Chromatography and PRM mass spectrometry, Easy-nLC II and Q exactive Orbitrap”.

- Antibodies: all monoclonal antibodies used in the immuno-affinity enrichment can be purchased from Cell Signaling Technology. See “SOP_Peptide immuno-affinity enrichment for the characterization of linearity and sensitivity of immuno-PRM assays using response curves” for details.

- Matrix: The background matrix is a tryptic digest of SKBR3 cells.

Solutions

- 20X Phosphate Buffered Saline (PBS) (CST, #9808)

- 10X Immuno-Affinity Purification (IAP) buffer (CST, #9993)

- A heavy-SIS mix with a concentration of 1 pmol/µL per peptide is prepared from the individual heavy peptide stocks using 5 % ACN-0.1 % TFA as a diluent.

- A light-SIS mix with a concentration of 1 pmol/µL per peptide is prepared from the individual light peptide stocks using 5 % ACN-0.1 % TFA as a diluent.

# Procedure

A. Preparation of antibody-protein A beads and validation samples

- Vortex to mix 50 % Protein A agarose bead slurry. Transfer 920 µL of bead slurry to a microcentrifuge tube.

- Centrifuge at 2000 × g for 20 sec and remove the supernatant.

- Add 1X PBS buffer to the beads, mix by inverting the tube 5 times, centrifuge for 30sec, and remove the supernatant

- Repeat step three for a total of 4 times PBS buffer washes.

- All subsequent wash steps are at 0-4 ˚C. In all wash steps, the supernatant should be removed reasonably well. Avoid removing the last few microliters, except in the last step, since this may cause inadvertent carry-over of the beads.

- The last time, after removing the supernatant, resuspend the beads in 1X PBS.

- Mix 230 µg of each antibody with beads. Pipet antibody solutions on top of the beads at the bottom of the tube to ensure immediate mixing. Avoid creating bubbles upon pipetting.

- Fill the microfuge tube with 1X PBS and seal it with parafilm in order to avoid leakage. Incubate overnight on a rotator at 4 ˚C.

- Wash beads with 1X PBS 4 times. Add 1X IAP buffer to reach a total volume of 920 µL.

- Aliquot antibody-bead slurry to 45 microcentrifuge tubes. Each tube contains 20 µL of antibody-bead slurry.  Store aliquots in 4 ˚C.

- Preparation of validation samples for determination of repeatability

1.  Preparation of background matrix and heavy peptides

- Based on preliminary reverse curve results, by following the CPTAC assay characterization guidance document and also considering the multiplex assays, heavy peptides are spiked into 100 µg of the digested lysate at three concentrations: high, medium, and low

    - 500 fmol/100 µg

    - 20 fmol/100 µg

    - 0.8 fmol/100 µg

- Following are specific preparation procedures for the mixture of background matrix spiked-in with heavy-SIS at three concentrations for all samples analyzed over 5 days.

- Prepare heavy-SIS peptides working solutions from 1 pmol/ µL stock solutions

  - High 500 fmol/µL

  - Medium 20 fmol/µL

  - Low 0.8 fmol/µL

- Then prepare the background matrix and heavy peptide mix

  - Spike 19 µL of peptide solution a (high) in 1.9 mL of 1 mg/ml SKBR3 tryptic digest, then aliquot 101 µL to each microcentrifuge tube for a total of 18 tubes

  - Spike 19 µL of peptide solution b (medium) in 1.9 mL of 1 mg/mL SKBR3 tryptic digest, then aliquot 101 µL to each microcentrifuge tube for a total of 18 tubes

  - Spike 19 µL of peptide solution c (low) in 1.9 mL of 1 mg/mL SKBR3 tryptic digest, then aliquot 101 µL to each microcentrifuge tube for a total of 18 tubes

- All samples are frozen at -80 ˚C

  - Peptide immuno-affinity enrichment

    - On each of the 5 days, remove and thaw three replicates of previously prepared samples at three concentrations.

    - High 500 fmol/100 µg

    - Medium 20 fmol/100 µg

    - Low 0.8 fmol/100 µg

    - Spike freshly prepared 10 µL of 1 fmol/µL light-SIS into each of 9 sample tubes

    - Transfer 20 µL of antibody-bead slurry to each of the sample tubes

    - Follow the steps of “immuno-affinity enrichment IV-XIV” in “SOP_ Peptide immuno-affinity enrichment for the characterization of linearity and sensitivity of immuno-PRM assays using response curves”.

| Run order | Sample |
| --- | --- |
| 1 | Synthetic light peptide mix |
| 2 | Wash |
| 3 | Wash |
| 4 | Solvent blank |
| 5 | Solvent blank |
| 6 | Low sample, repeat 1 |
| 7 | Low sample, repeat 2 |
| 8 | Low sample, repeat 3 |
| 9 | Wash |
| 10 | Medium sample, repeat 1 |
| 11 | Medium sample, repeat 2 |
| 12 | Medium sample, repeat 3 |
| 13 | Wash |
| 14 | High sample, repeat 1 |
| 15 | High sample, repeat 2 |
| 16 | High sample, repeat 3 |
| 17 | Wash |
| 18 | Wash |
| 19 | Solvent blank |
| 20 | Solvent blank |

  - Stage tip cleanup

Perform stage tip C18 cleanup of peptides according to the SOP of “SOP_ Peptide immuno-affinity enrichment for the characterization of linearity and sensitivity of immuno-PRM assays using response curves”

- Analysis by PRM mass spectrometry

  - Resuspend dried samples in 10 µL of 5 % ACN-0.1 % TFA reconstitution buffer, sonicate for 1 min, and briefly centrifuge the sample tubes to remove the evaporated solvent. Set up parameters according to the SOP of “Liquid Chromatography and PRM mass spectrometry, Easy-nLC II and Q Exactive Orbitrap.”

  - Sample run order on the LC-MS system

The repeatability samples are run in increasing order of heavy peptide level, with one injection per capture replicate. Light-SIS mix (1 fmol) is injected to determine the actual retention times of the peptides. This is followed by extensive washes. Solvent blanks (0.1 % FA) are used to confirm the removal of peptides on-column. The table below illustrates the run order.

  - Data analysis

Data analysis is done by open-source software Skyline. Raw files are imported into Skyline, and extracted ion chromatograms of all transition ions are integrated.  Manually check all peaks of light and heavy peptides to confirm correct integration. https://skyline.ms/project/home/software/skyline/begin.view