# STANDARD OPERATING PROCEDURE

Title: Mini-Validation of Repeatability. Version #: 1; Author: Hui Zhang Lab. Date: 05/08/2014.

# Purpose

The purpose of this document is to describe the characterization of a set of assays based on its repeatability of measurement over 5 days. This is to estimate the performance of the assay measured in a complex sample across several days.

# Scope

This procedure addresses the preparation and running of samples for generating the validation samples in accordance with CPTAC Assay Characterization Guidance Experiment #2.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures.  The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

- Microcentrifuge

- Vacuum centrifuge

# Materials

- Water: Optima LC/MS-grade (Fisher Scientific; cat. # W6-4)

- Acetonitrile: Optima LC/MS-grade (Fisher Scientific; cat. # A955-4)

- Formic Acid: LC-MS Ultra (Sigma-Aldrich; cat. # 14265)

- Methanol: Optima LC/MS-grade (Fisher Scientific; cat. # A456-4)

- Ammonium formate (Sigma-Aldrich; cat. # 70221)

- Ammonium hydroxide (Sigma-Aldrich; cat. # 320145)

- Polysulfoethyl A TopTips 100 – 200 µL (Glygen; cat. # TT2SSA.96)

# Reagents

- 10 mM Ammonium formate in 25% ACN, pH 3.0

- 500 mM Ammonium formate in 25% ACN, pH 6.8

- 80:15:5 (vol:vol:vol) Methanol: Water: Ammonium hydroxide

- Stable isotope-labeled standards (SIS)

  - Crude SIS formerly N-glycosylated peptides were synthesized (~60% purity) by Thermo Fisher Scientific (PEPotec SRM peptide library). SIS peptides incorporated a fully atom labeled 13C and 15N isotope at the C-terminal lysine (K) or arginine (R) position of each tryptic peptide, resulting in a mass shift of +8 or +10 Da, respectively. Deamidated asparagine residues corresponding to N-glycosylation sites were synthesized as aspartic acid residues. Peptides were provided in 0.1% TFA/50% ACN and stored at -80 °C until use.

  - A stock SIS mix was prepared at a concentration of 1320 pmol/µL for each peptide, followed by de-salting via strong cation exchange (SCX) as indicated in Procedure 1, below. The peptide recovery following SCX clean-up was assumed to be 50%.

- Matrix

  - A background matrix consisting of trypsin digested serum-derived N-glycopeptides was prepared according to the SOP entitled “Trypsin digestion of serum and enrichment of N-glycopeptides using an automated liquid handler”. This background matrix was used for the preparation of the response curves and for the preparation of the mini-validation of repeatability experiments. Given the limitation of a maximum yield of 20 µg of glycopeptides per preparation of 40 µL serum, N-glycopeptides prepared using multiple DART pipette tips packed with hydrazide resin were combined to achieve the required amount of N-glycopeptides that were needed for the repeatability study.

# Procedure

- SCX de-salting of stock SIS mix

  - All centrifugation steps are performed at 2,000 rpm for 1.5 min, unless otherwise specified. The maximum binding capacity of each TopTip is 1 mg.

    - Condition Polysulfoethyl A TopTip 2x with 330 µL of Methanol

    - Wash 2x with 330 µL of 10 mM Ammonium Formate in 25% ACN, pH 3.0

    - Wash 2x with 330 µL of 500 mM Ammonium Formate in 25% ACN, pH 6.8

    - Wash 2x with 330 µL of 10 mM Ammonium Formate in 25% ACN, pH 3.0

    - Wash 2x with 330 µL of Water

    - Wash 4x with 330 µL of 10 mM Ammonium Formate in 25% ACN, pH 3.0

    - Slowly load acidified sample (pH < 3.0) 2x; centrifuge at 1,100 rpm for 5 min

    - Wash 6x with 330 µL of 10 mM Ammonium Formate in 25% ACN, pH 3.0

    - Allow TopTip to dry out. Elute sample 2x with 300 µL of 80:15:5 (vol:vol:vol) Methanol: Water: Ammonium Hydroxide; centrifuge at 1,100 rpm for 5 min

    - Dry eluted sample in a vacuum centrifuge

- Determination of spike levels and preparation of samples

  - Peptides were multiplexed according to the LLOQ and linear range determined from the response curves in the CPTAC Assay Characterization Guidance Experiment #1 in order to prepare Validation samples at an appropriate concentration. The three SIS spike levels were as follows:

    - Low: 0.288 pmol/µL (1.728 pmol on column)

    - Medium: 7.5 pmol/µL (45 pmol on column)

    - High: 12 pmol/µL (72 pmol on column)

These concentrations were selected to approximate 1.5-3.0x LLOQ (Low), 50-100x LLOQ (Medium), and >100x LLOQ (High).

  - Dilute matrix to 0.25 µg/µL with 2% ACN/0.2% formic acid.

  - Prepare dilution mix containing 366.67 µL matrix (0.25 µg/µL), 110 µL iRT standard diluted 1:10 and 623.33 µL 0.1 % formic acid (final concentration of matrix in dilution mix is 0.083 µg/µL and final dilution of iRT is 1:100). A total volume of dilution mix should be prepared to permit at least 15 injections of each sample.

  - Prepare SIS stock by taking 84 µL of SIS (660 pmol/µL) and drying in a vacuum centrifuge. Resuspend in 369.6 µL dilution mix (prepared in Step 1c, above) for a final SIS stock concentration of 150 pmol/µL.

  - Prepare “High” concentration sample by adding 9.6 µL of SIS to 110.4 µL of dilution mix.

  - Prepare “Medium” concentration sample by adding 6.0 µL of SIS to 114 µL of dilution mix.

  - Prepare “Low” concentration sample by adding 2.3 µL of 1:10 dilution of SIS to 117.7 µL of dilution mix.

  - Store samples at -80 °C until LC-PRM MS analysis.

- Execution of LC-PRM MS analysis

  - Vortex samples, centrifuge briefly and transfer to autosampler vials. Add sufficient volume to each vial for all of the replicate injections.

  - Perform LC-PRM MS analysis according to the SOPs entitled “Liquid Chromatography, Dionex UltiMate 3000 RSLCnano LC System” and “Parallel Reaction Monitoring (PRM) Mass Spectrometry, Q-Exactive”.

- Run order

  - To avoid artificially minimizing variability, the run order of the samples should be randomized. To minimize carryover, 1 wash is inserted after the “Low” and “Medium” samples, and 2 washes are inserted after the “High” samples.

# Referenced Documents

- SOP Liquid Chromatography, Dionex UltiMate 3000 RSLCnano LC System

- SOP Parallel Reaction Monitoring (PRM) Mass Spectrometry, Q-Exactive

- SOP Trypsin digestion of serum and enrichment of N-glycopeptides using an automated liquid

handler